## Summary
2 Channel Dual Voltage-Current Meter: firmware v1.0.0 release report
First production release. Compared against v0.3.0. Updated 27 Sep 2026 for the final release file (voltage calibration line).
Release
Firmware file | 2_channel_dual_voltage_current_meter_fw_v1.0.0.ino (final)
Target | Arduino Nano (ATmega328P), Arduino AVR core
Tested binary | The final release file compiles to exactly the build every test suite ran on (production and debug)
Voltage calibration | displayed = measured x 0.99578 + 31 mV, rounded to 10 mV: 0.790-20.110 V, then OL
Verdict | Ready for release: no failures in any simulated test; bench voltage check agrees with a Fluke 15B+
Release condition | Release after the remaining Bench Test Plan checks pass
Simulated tests (v1.0.0)
Test cases | 101
Passed | 98
Failed | 0
Observations | 3 | Behavior kept from v0.3.0 on purpose; see Findings.
Unit tests (functions called directly on the chip)
Unit test cases | 234322
Mismatches against the models | 0 | Must be 0.
Parity with v0.3.0 (every ADC code 0-1023)
Voltage: identical | 595
Voltage: within 10 mV (calibration line) | 423 | The line replaces the 7 offset bands; mean difference 0.0 mV.
Voltage: -20 mV at the end of the old +30 mV band | 3 | 4.13-4.17 V: the old band over-corrected there.
Voltage: low range fixed (0.82, 0.84 V) | 2 | v0.3.0 showed 0.790 / 0.800 here and ran backwards.
Voltage: ADC full scale now OL | 1
Current: identical | 852
Current: +0.01 A (nearest step) | 58 | v1.0.0 picks the closer table value; v0.3.0 leaned to the lower one.
Current: gap bug fixed (2.49/2.50 A) | 2
Current: above 5.00 A now OL | 112
Unexplained differences | 0 | Must be 0.
Findings
Open: check on bench | 4
Open: recommendations | 8
Fixed in the final file | 2 | F-02 and F-03: removed by the calibration line.
Stability (display changes in 1,500 noisy cycles)
v0.3.0 | 528
v1.0.0 | 4
Reduction | 99.24%
Voltage accuracy
Estimated spec | ±(0.1 % + 2 digits) | Error budget on the Voltage Accuracy sheet; 1 digit = 10 mV.
Fluke 15B+ points inside its tolerance | 10 of 10
Largest difference where the Fluke can decide (mV) | #VALUE! | Below 4 V; above 4 V a better reference is needed (F-13, BT-49).
Keysight E36312A vs this meter at 12 V | 0.24 | About 4x more accurate; Siglent SPD3303X-E is about 2.5x less accurate.
Memory and timing (v1.0.0 production build)
Flash | 9358 | bytes
Free RAM at peak | 1565 | bytes
Free RAM at peak, v0.3.0 | 596 | bytes
Measurement work per 560 ms cycle | 42.7 | ms (typical)
Worst-case measurement work | 56.604 | ms (samples in worst order for sorting)
Bench test plan (fill in on the hardware)
Passed | 0
Failed | 0
Not done yet | 49
Before flashing a production unit
1. Confirm DEBUG_FLAG is 0 in the file you flash.
2. Check the board uploads with Processor "ATmega328P" (Optiboot). Old bootloader + watchdog can reset endlessly.
3. Keep VOLTAGE_CAL_GAIN 0.99578 and VOLTAGE_CAL_OFFSET_MV 31.0 unless BT-49 with a better reference shows otherwise.
4. Run the Bench Test Plan, including BT-43 (outlier rejection on your real load) and BT-49 (voltage 5-20 V).
Sheets in this workbook
Changes | Every modification from v0.3.0 to v1.0.0
Enhancements | Improvements with measured numbers
Test Cases | 101 test cases with v0.3.0 and v1.0.0 results
Unit Tests | Function-by-function results
ADC Parity | All 1,024 ADC codes, voltage, current, LED and LCD row
Stability Tests | Noise, dropout and ramp tests; outlier rejection limits
LCD Scenarios | Exact LCD rows in over range, LED, alert and crosstalk events
Findings | Open and closed findings for v1.0.0
Memory & Timing | Flash, RAM, stack, cycle time, watchdog margin
Bench Test Plan | Hardware checks to fill in
Voltage Accuracy | Fluke 15B+ bench check, error budget, comparison with commercial supplies
Method | How everything was tested

## Changes
Modifications from v0.3.0 to v1.0.0
v0.3.0 = 2_channel_dual_voltage_current_meter_fw_v0.3.0.ino. v1.0.0 = 2_channel_dual_voltage_current_meter_fw_v1.0.0.ino.
ID | Area | Type | v0.3.0 behavior | v1.0.0 behavior | User visible | Verified by
CH-01 | Current | Bug fix | Linear search in two halves; 2.219-2.228 V (2.49-2.50 A) fell in neither half | Binary search over the whole table | Yes | C-03
CH-02 | Current | Behavior change | Picks the lower step unless the reading is > 5 mV above it | Picks the truly nearest step (58 of 912 codes show +0.01 A) | Yes | C-02, C-05
CH-03 | Current | Bug fix | Above 5.00 A returns 999.23: display "999.2A0", LED off | "OL" in the current field, LED on | Yes | C-04, L-03
CH-04 | Voltage | Enhancement | ADC full scale shows 20.130 V | "OL" in the voltage field | Yes | V-02
CH-05 | ADC | Enhancement | 1 sample per reading | 64 samples per channel (16 for temperature), fractional average | Yes (stability) | V-06 to V-08, C-07, C-08
CH-06 | ADC | Enhancement | None | Outlier rejection: samples > 8 counts from the median are ignored | Yes (dropouts) | A-07 to A-09
CH-07 | ADC | Enhancement | 1 settling read after a channel switch | 2 settling reads | No | A-03, A-06
CH-08 | ADC | Enhancement | Fixed sample timing | 0-7 us random jitter between samples | No | A-04
CH-09 | Display | Enhancement | None | Hysteresis: 3 mV on voltage steps, 1.5 mV on current steps | Yes (no flicker) | V-07, V-09, C-07
CH-10 | Display | Enhancement | n/a | Hysteresis cleared at startup, over range, below 0.790 V and on alert changes | No | T-05
CH-11 | Temperature | Enhancement | Alert on and off at 0.680 V (toggles near threshold) | On at 0.680 V, off below 0.670 V | Yes | T-02, T-03
CH-12 | Display | Bug fix | Alert screen cleared every cycle (flicker); leftovers "-T"/"LE" after the alert | Alert drawn once; screen cleared once on exit | Yes | D-05, D-06
CH-13 | Display | Behavior change | Alert label "IC6-TEMP" | Alert label "VREF-TEMP" | Yes | D-05
CH-14 | Display | Bug fix | Fields printed in pieces; stale "0" in column 15 after an error value | Each row built as a full 16-character line | Yes | D-01, D-02
CH-15 | Display | Behavior change | 4 decimals below 10 V ("5.0300V"), 3 above ("12.030V") | Always 3 decimals, right-aligned (" 5.030V") | Yes | D-04
CH-16 | Timing | Enhancement | delay(10) + delay(550): real period about 573 ms | millis() scheduler: exactly 560 ms, rollover safe | No | S-01 to S-03
CH-17 | Reliability | Enhancement | None | 2 s watchdog; reset cause captured; watchdog disabled early after a reset | Yes (splash message) | W-01, W-02
CH-18 | Startup | Enhancement | Analog pins INPUT only | DIDR0 buffers off, unused pins pulled up, D13 and LEDs LOW, reference selected first | No | I-01 to I-06
CH-19 | Startup | Bug fix | pinMode(A6) outside the pin tables | Removed | No | I-07
CH-20 | Memory | Enhancement | Lookup table in RAM (1,319 B static RAM) | Table in flash (PROGMEM), strings in flash (F()): 273 B static RAM | No | M-01, M-02
CH-21 | Serial | Enhancement | Serial.begin(9600) with no output | No serial in production; debug build at 115200 with per-cycle diagnostics | No | G-01
CH-22 | Code | Refactor | Globals per channel, magic numbers, 999.23 sentinel | Channel struct, Reading/ReadingStatus, named constexpr constants, prototypes, static_asserts | No | B-01, B-04
CH-23 | Startup | Refactor | lcd.begin() called twice | Called once; splash text unchanged | No | I-08
CH-24 | Voltage | Calibration | 7 offset bands (+30 to -40 mV) added after truncating to 10 mV; special low range (0.790 / 0.810 V below 0.800 V) | One calibration line: displayed = measured x 0.99578 + 31 mV (fit of the same bands), rounded to the nearest 10 mV; 0.000 below 0.790 V. Constants VOLTAGE_CAL_GAIN / VOLTAGE_CAL_OFFSET_MV (1.0 / 0.0 = no compensation). 410 B less flash. | Yes (within 10 mV; no backwards steps) | V-01, V-04, V-05, V-09, V-11, V-12

## Enhancements
What v1.0.0 improves, with measured evidence
All figures are measured in the simulation runs (see Method). Improvement columns are formulas.
Enhancement | Metric | v0.3.0 | v1.0.0 | Improvement | Evidence
Stable voltage display | Display changes, 10 mV boundary (ADC 611.40), 300 cycles | 158 | 3 | 98.1% | V-07
Stable voltage display | Display changes inside a step (ADC 611.15), 300 cycles | 188 | 0 | 100% | V-08
Stable current display | Display changes, 0.01 A boundary, 300 cycles | 142 | 1 | 99.3% | C-07
Stable current display | Display changes inside a step, 300 cycles | 30 | 0 | 100% | C-08
Dropout immunity | Cycles showing 0.02 A with 2% dropouts, of 300 | 5 | 0 | 100% | A-07
Voltage resolution | Distinct displayed voltages reachable | 984 | 1934 | 1.9654 | V-06, V-11; fine sweep
Correct current in 2.49-2.50 A | ADC codes showing an error value | 2 | 0 | 100% | C-03
Clear over-range display | Codes showing garbage "999.2A0" | 112 | 0 | 100% | C-04
Clean LCD rows | Sweep steps with a leftover character | 569 | 0 | 100% | D-01
No alert flicker | Screen clears per alert cycle | 1 | 0 | 100% | D-05
Alert hysteresis | Alert state changes while hovering 0.684/0.679 V after the alert is on (3 cycles) | 3 | 0 | 100% | T-03
Crosstalk margin | Contaminated reads tolerated after a channel switch | 1 | 2 | 2 | A-05, A-06
RAM headroom | Free RAM at peak stack (bytes) | 596 | 1565 | 2.6258 | M-02
Timing precision | Measurement period (ms) | 573 | 560 | Exactly 560 ms (was 573 ms) | S-01, S-03
Self-recovery | Time to recover from a firmware hang (ms) | Never (hangs) | 2029 | Recovers automatically | W-01
Smooth voltage calibration | Places where a rising voltage makes the display step back (fine sweep) | 2 | 0 | 100% | V-09, F-02, F-03
Improvement = 1 - v1.0.0 / v0.3.0 for "fewer" rows, v1.0.0 / v0.3.0 for "x" rows.

## Test Cases
Test cases and results
Results for v1.0.0 and v0.3.0. Source: simulated ATmega328P test runs (see Method).
ID | Group | Test | Stimulus | Expected (v1.0.0) | v0.3.0 result | v1.0.0 result | Status | Method
B-01 | Build | Production build compiles (DEBUG_FLAG 0) | avr-gcc 7.3.0, strict C++ (no IDE auto-prototypes), -Wall -Wextra | No errors, no warnings | Compiles only with IDE auto-prototypes | No errors, no warnings. 9,358 B flash, 273 B SRAM | PASS | Compiler
B-02 | Build | Debug build compiles (DEBUG_FLAG 1) | Same, DEBUG_FLAG 1 | No errors, no warnings | n/a (no debug mode) | No errors, no warnings. 10,358 B flash, 287 B SRAM | PASS | Compiler
B-03 | Build | Release file is set to production | Check DEBUG_FLAG in the file to be flashed | DEBUG_FLAG 0 | n/a | DEBUG_FLAG 0 | PASS | Code review
B-04 | Build | Compile-time checks | static_assert: table size 501, alert OFF < ON, sample counts <= 64 | All hold | None present | All hold | PASS | Compiler
I-01 | Startup | External reference selected before first ADC read | Record order of analogReference / analogRead | analogReference(EXTERNAL) before any analogRead | OK (set in setup, reads in loop) | OK (first line of setup) | PASS | Simulation
I-02 | Startup | Analog inputs A0-A3 configured | Record pinMode calls | INPUT, no pull-up | INPUT | INPUT | PASS | Simulation
I-03 | Startup | Digital input buffers off on A0-A3 | Read DIDR0 after setup | DIDR0 = 0x0F | DIDR0 = 0x00 | DIDR0 = 0x0F | PASS | Simulation
I-04 | Startup | LED pins D3/D5 | Record pinMode / digitalWrite | OUTPUT, LOW at start | OUTPUT (default LOW) | OUTPUT, explicitly LOW | PASS | Simulation
I-05 | Startup | Unused pins D2, D4, D6, A4, A5 | Record pinMode | INPUT_PULLUP (not floating) | Not configured (floating) | INPUT_PULLUP | PASS | Simulation
I-06 | Startup | On-board LED D13 | Record pinMode / digitalWrite | OUTPUT, LOW | Not configured | OUTPUT, LOW | PASS | Simulation
I-07 | Startup | No pinMode() on A6 (pin 20) | Record pinMode calls; inspect pin tables in the binary | Not called (A6 is analog-only, outside Arduino pin tables) | Called: reads past pin table (no effect in this build) | Not called | PASS | Simulation + binary inspection
I-08 | Startup | Splash screen | Capture LCD during the 2 s splash | "STARTING" / "PLEASE WAIT" | ¦ STARTING ¦ PLEASE WAIT ¦ | ¦ STARTING ¦ PLEASE WAIT ¦ | PASS | Simulation
I-09 | Startup | Splash after a watchdog reset | Force a hang, capture splash on reboot | "WATCHDOG RESET" on row 2 | n/a (no watchdog) | ¦ STARTING ¦ WATCHDOG RESET ¦ | PASS | Simulation
A-01 | ADC | Conversions per measurement cycle | Count analogRead calls in one cycle | 282 = 18 temperature + 4 x 66 | 10 (2 per pin) | 282 (A6 18, A0-A3 66 each) | PASS | Simulation
A-02 | ADC | Channel read order | Record pin sequence | A6, A3, A2, A0, A1 (same as v0.3.0) | A6, A3, A2, A0, A1 | A6, A3, A2, A0, A1 | PASS | Simulation
A-03 | ADC | Settling reads after channel switch | Count discarded reads | 2 per switch | 1 per switch | 2 per switch | PASS | Simulation
A-04 | ADC | Sample jitter distribution | Histogram of delayMicroseconds() arguments over the full run | Uniform 0-7 us | n/a | 0-7 us, each value 12.5% (219,884-219,971 per bin) | PASS | Simulation
A-05 | ADC | Crosstalk: 1 contaminated read after switch | Model: first read after a switch returns previous channel value | Correct reading | Correct | Correct | PASS | Simulation (model)
A-06 | ADC | Crosstalk: 2 contaminated reads after switch | Model: first 2 reads return previous channel value | Correct reading | Wrong: shows 0.0000V / 0.000A | Correct: 17.700V / 0.720A | PASS | Simulation (model)
A-07 | ADC | Single-sample dropouts (original symptom) | 0.72 A, ±0.5 count noise, 2% of samples drop to 0.02 A, 300 cycles | Steady 0.720A | 0.02A shown in 5 of 300 cycles | 0.720A in all 300 cycles (with outlier rejection off, OUTLIER_LIMIT_COUNTS = 1023: 215 display changes, 0.65-0.72 A) | PASS | Simulation
A-08 | ADC | Outlier rejection disabled (limit 1023) | Full test run with OUTLIER_LIMIT_COUNTS = 1023 | Identical to a plain average of all samples | n/a | 6,473 of 6,473 rows identical to a plain-average build of the final file | PASS | Simulation
A-09 | ADC | Outlier rejection is neutral for normal signals | All sweeps, ramps, noise, OL, LED, temperature, crosstalk tests | Identical to plain average | n/a | All rows identical except the 300-cycle dropout test (A-07) | PASS | Simulation
V-01 | Voltage | Full ADC sweep, codes 0-1022, parity with v0.3.0 | Each code, fresh state, both channels | Within 10 mV of v0.3.0 except the intended fixes | Reference | 595 identical, 423 within ±10 mV (line instead of bands), 3 at -20 mV (4.13-4.17 V, end of the old +30 mV band), 2 low-range fixes (codes 40-41), 1 OL | PASS | Simulation
V-02 | Voltage | ADC full scale (code 1023) | Voltage ADC saturated | "OL" | 20.130V (not a real value) | OL | PASS | Simulation
V-03 | Voltage | Independent model check | Python float32 re-implementation of both algorithms (v0.3.0 offset bands, v1.0.0 calibration line) | Both firmwares match their model | Matches 1024/1024 | Matches 1023/1023 codes, plus all 51,151 points of a fine sweep (ADC 0-1023 in 0.02 steps) | PASS | Simulation + model
V-04 | Voltage | Low range near 0.790 V | Codes 0-45 | 0.000 below 0.790 V (after calibration), then only rising values | Codes 38-41: 0.000, 0.790, 0.790, 0.800; runs backwards at 0.800 V (F-02) | Codes 38-42: 0.000, 0.800, 0.820, 0.840, 0.860; always rising | PASS | Simulation
V-05 | Voltage | Calibration line | VOLTAGE_CAL_GAIN 0.99578, VOLTAGE_CAL_OFFSET_MV 31.0, whole range | Smooth correction: +27 mV at 1 V to -53 mV at 20 V, no band steps | 7 offset bands, +30 to -40 mV | Matches the model at every point. Versus v0.3.0: mean 0.0 mV, 99.7% of fine-sweep points within ±10 mV | PASS | Simulation
V-06 | Voltage | Resolution on a slow ramp | Codes 200-260 (3.9-5.1 V) in 0.1 steps, up and down | All 118 10 mV values | 59 changes (skips every other 10 mV value) | 117 changes (every 10 mV value, 3.960-5.130 V); never more than 10 mV from ideal | PASS | Simulation
V-07 | Voltage | Stability on a 10 mV boundary | ADC 611.40 (12.035 V after calibration, exactly on a step boundary), ±1 count noise, 300 cycles | Few display changes | 158 changes (12.01/12.03/12.05) | 3 changes (12.03/12.04) | PASS | Simulation
V-08 | Voltage | Stability inside a step | ADC 611.15 (12.030 V), ±1 count noise, 300 cycles | No display changes | 188 changes | 0 changes | PASS | Simulation
V-09 | Voltage | Hysteresis correctness | 2,886 ramp cycles (simulation), 40,922-step function ramp ADC 0-1023 up and down, 96,962 unit cases | Never more than 1 step from nearest value; never moves against the input | n/a | Max 1 step (0.01 V); 0 reversals anywhere (F-02 and F-03 closed) | PASS | Simulation + model
V-10 | Voltage | Channel 2 identical to channel 1 | Same input on both channels, all sweeps | Identical rows | Identical | Identical | PASS | Simulation
C-01 | Current | Full ADC sweep, parity with v0.3.0 | Codes 0-1023, fresh state | Same as v0.3.0 except intended fixes | Reference | 852 identical, 58 +0.01 A, 2 gap fixed, 112 OL | PASS | Simulation
C-02 | Current | Rounding to nearest 0.01 A step | 58 codes where v0.3.0 rounds down | Nearest table step | Lower step (5 mV rule) | Nearest step (+0.01 A vs v0.3.0) | PASS | Simulation + model
C-03 | Current | 2.49-2.50 A gap | Codes 455 and 456 | 2.490A and 2.500A | 999.2A0 (error value) | 2.490A and 2.500A | PASS | Simulation
C-04 | Current | Above 5.00 A | Codes 912-1023 | "OL" | 999.2A0 (garbage), LED off | OL, LED on | PASS | Simulation
C-05 | Current | Independent model check | Python nearest-step model | Match | n/a | 1024 of 1024 | PASS | Model
C-06 | Current | Slow ramp 0.49-0.78 A | Codes 100-160 in 0.1 steps, up and down | Monotonic, max 1 step from nearest | Monotonic, max 0.01 A | Monotonic, max 0.01 A | PASS | Simulation
C-07 | Current | Stability on a 0.01 A boundary | 0.71/0.72 A boundary, ±1 count noise, 300 cycles | Few display changes | 142 changes | 1 change | PASS | Simulation
C-08 | Current | Stability inside a step | 0.72 A, ±1 count noise, 300 cycles | No display changes | 30 changes (0.72/0.73) | 0 changes | PASS | Simulation
C-09 | Current | Channel 2 identical to channel 1 | Same input on both channels | Identical rows | Identical | Identical | PASS | Simulation
D-01 | Display | No leftover characters in normal rows | Current sweep 0-1023 | Column 15 always blank | "0" left in column 15 in 569 of 1024 steps | Blank in 1024 of 1024 | PASS | Simulation
D-02 | Display | Return from current over range | OL then 0.72 A | "0.720A" clean | "0.720A0" | "0.720A" | PASS | Simulation
D-03 | Display | Voltage over range | ADC 1023 | "OL" in voltage field | "20.130V" | " OL" | PASS | Simulation
D-04 | Display | Row format | All tests | 16 characters, 3 decimals | "5.0300V" below 10 V, "12.030V" above | " 5.030V" / "12.030V", always 3 decimals | PASS | Simulation
D-05 | Display | Temperature alert screen | Temperature 0.684 V | Alert text, drawn once | "IC6-TEMP" / "ALERT", screen cleared every cycle (flicker) | "VREF-TEMP" (column 4, as v0.3.0) / "ALERT", cleared once on entry | PASS | Simulation
D-06 | Display | Screen after the alert ends | Temperature back to 0.645 V | Clean measurement rows | Leftover "-T" and "LE" from alert text | Clean (cleared once on exit) | PASS | Simulation
L-01 | LED | LED at 0.00 A | Current code 0 and 1 | Off | Off | Off | PASS | Simulation
L-02 | LED | LED at 0.01-5.00 A | Current codes 2-911 | On | On | On | PASS | Simulation
L-03 | LED | LED above 5.00 A | Current codes 912-1023 | On (current is flowing) | Off | On | PASS | Simulation
L-04 | LED | LEDs during temperature alert | Alert active | Documented behavior | Keep last state | Keep last state (unchanged from v0.3.0) | OBSERVATION | Simulation
T-01 | Temperature | Alert turns on | Temperature code 140 (0.684 V) | Alert on at >= 0.680 V | On | On | PASS | Simulation
T-02 | Temperature | Alert hysteresis | Codes 139, 138 after alert | Stays on until < 0.670 V | Turns off at 139 (0.679 V) | Stays on at 139/138; off at 137 (0.669 V) | PASS | Simulation
T-03 | Temperature | Hovering at the threshold | Alternating codes 140/139 | No toggling | Toggles every cycle | Stays in alert | PASS | Simulation
T-04 | Temperature | Measurements pause during alert | Count ADC reads during alert | Only temperature is read | Channels read but not shown | 18 reads (temperature only) | PASS | Simulation
T-05 | Temperature | Hysteresis state reset after alert | Alert on then off | Previous steps cleared | n/a | Cleared (NO_STEP) | PASS | Code review + simulation
S-01 | Timing | Refresh period | loop() at +100, +559, +560 ms | Measure only at +560 ms | n/a (delay based) | 0, 0, 282 reads | PASS | Simulation
S-02 | Timing | millis() rollover (49.7 days) | Last measurement 0xFFFFFE00, now 0xFFFFFFF0 then 0x00000030 | No measurement at 496 ms, measurement at 560 ms | n/a | Correct | PASS | Simulation
S-03 | Timing | Measurement cycle work time | Real firmware, real ADC timing and LCD library | Well under 560 ms | 13 ms work + 560 ms delays = 573 ms period | 42.7 ms work, 560 ms period (debug build 47.2 ms) | PASS | Simulation (real firmware)
W-01 | Watchdog | Recovery from a firmware hang | Infinite loop inside analogRead(A2) | Reset within ~2 s, normal operation after | n/a (hangs forever) | Reset after 2,029 ms; reset flags 0x08; normal readings after reboot | PASS | Simulation
W-02 | Watchdog | No spurious resets | Full test run, about 6,500 cycles | No unexpected reset | n/a | None | PASS | Simulation
M-01 | Memory | Static RAM | Compiler report | Low | 1,319 B (64%) | 273 B (13%) | PASS | Compiler
M-02 | Memory | Peak stack | RAM painting, real firmware | Enough free RAM | 133 B, 596 B free | 210 B, 1,565 B free | PASS | Simulation (real firmware)
B-05 | Build | Release file compiles to the tested binary | Compare compiled binaries byte by byte | Identical binaries | n/a | Production and debug builds of the final file are exactly the binaries every suite ran on | PASS | Compiler
B-06 | Build | Extra compiler warnings | -Wshadow -Wpedantic -Wcast-qual -Wdouble-promotion -Wundef -Wuninitialized -Wlogical-op -Wduplicated-cond | No real issues | n/a | Only -Wdouble-promotion on fabs() (lines 492, 533): double = float on AVR, no effect (F-09) | PASS | Compiler
B-07 | Build | Static analysis | cppcheck 2.13 --enable=all --inconclusive, debug build | No defects | n/a | No defects (only "setup/loop/captureResetFlags never used": called by the Arduino core and startup code) | PASS | Static analysis
B-08 | Build | Compile-time checks block bad settings | Break each check on purpose: 65 samples, 0 samples, alert OFF > ON, table with 500 entries | Build fails with a clear message | n/a | All 4 builds fail with their own message ("Too many samples", "At least one sample", ...) | PASS | Compiler
R-01 | Regression | Full system bench on the final release file | Same 6,473-row run, compared with the previous v1.0.0 build (7-band voltage table) | Only the voltage values change, as intended | n/a | Current, LEDs, screen clears and ADC read counts identical in all 6,473 rows. Voltage differs only where the calibration line changes it; all 2,048 sweep values match the model. Alert text at column 4, as v0.3.0 | PASS | Simulation
R-02 | Regression | Unit, edge and timing suites on the final file | Unit (125,583 lines), edge (3,445 lines), fine voltage sweep and ramps, timing on the real firmware | All pass | n/a | 0 mismatches in every suite; timing 42.7 ms per cycle, peak stack 210 B | PASS | Simulation
U-01 | Unit: averaging | All samples equal | All 0, all 1023, all 500 | Exact value | n/a | 0.0000, 1023.0000, 500.0000 | PASS | Unit test
U-02 | Unit: averaging | Sorting | Ascending 0-63 and descending 63-0 | Same result | n/a | 32.0000 both | PASS | Unit test
U-03 | Unit: averaging | Dropouts and spikes rejected | 62x131 + 2x4; 60x131 + 4x1023 | 131 | n/a | 131.0000 both | PASS | Unit test
U-04 | Unit: averaging | Rejection limit exactly 8 counts | 60x500 + 2x508 + 2x509; 60x500 + 2x492 + 2x491 | 508/492 kept, 509/491 dropped | n/a | 500.2581 and 499.7420 (exact expected values) | PASS | Unit test
U-05 | Unit: averaging | Two levels (load step inside the window) | 32x100 + 32x200; 33x100 + 31x200 | Majority level | n/a | 200.0000 and 100.0000 | PASS | Unit test
U-06 | Unit: averaging | Normal noise equals plain mean | 64 random samples 128-135 | Plain mean 131.65625 | n/a | 131.6562 | PASS | Unit test
U-07 | Unit: averaging | Temperature (16 samples) | 14x139 + 2x0 | 139 | n/a | 139.0000, 18 reads | PASS | Unit test
U-08 | Unit: averaging | Sample count above 64 is capped | Request 100 samples | 64 used, no array overflow | n/a | 300.0000, 66 reads (2 settling + 64) | PASS | Unit test
U-09 | Unit: averaging | Sample count 1 and 0 | Request 1 and 0 samples | 1: the sample; 0: not used by firmware | n/a | 1: 777.0000. 0 is blocked at compile time (static_assert "At least one sample") | PASS | Unit test
U-10 | Unit: voltage | computeVoltage() at every step boundary | Steps 79-2017 (0.785-20.165 V) x 10 offsets (-3.1 to +5 mV) x 5 previous states, plus OL and 0.790 V threshold edges | Matches float32 model | n/a | 96,962 cases, 0 mismatches | PASS | Unit test + model
U-11 | Unit: current | computeCurrent() at every table midpoint and entry | 500 intervals x 10 offsets (±1.6 mV) x 5 previous states + OL edges | Matches float32 model | n/a | 25,001 cases, 0 mismatches | PASS | Unit test + model
U-12 | Unit: current | findNearestCurrentStep() | 0 to 4,451 mV in 0.25 mV steps | Nearest table entry | n/a | 17,805 cases, 0 mismatches | PASS | Unit test + model
U-13 | Unit: display | formatChannelRow() for every reachable value | 0.000 and all 1,933 values 0.790-20.110 V, all 501 current values, 3 OL combinations | Exact 16-character rows, terminated | n/a | 2,438 rows, 0 mismatches | PASS | Unit test + model
U-14 | Unit: voltage | 0.790 V display threshold | ADC 38.600 / 38.650 / 38.658 / 38.660 / 38.670 / 38.700 | 0.000 below 790 mV after calibration (ADC 38.659), 0.790 from there | n/a | 0.000, 0.000, 0.000, 0.790, 0.790, 0.790 | PASS | Unit test + model
U-15 | Unit: voltage | Rounding to the nearest 10 mV | Every step boundary ±0.01 mV, no previous state | Halfway values round up (v0.3.0 truncated) | n/a | Part of U-10: 0 mismatches | PASS | Unit test + model
U-16 | Unit: temperature | Alert logic at the thresholds | 0.669999 / 0.67 / 0.679999 / 0.68 V from both states | On at 0.680, off below 0.670 | n/a | 16 cases, 0 mismatches | PASS | Unit test
U-17 | Unit: hysteresis | isAdjacentStep() | NO_STEP, equal, ±1, ±2, extremes | Adjacent only for ±1 | n/a | 12 cases, 0 mismatches | PASS | Unit test
U-18 | Unit: hysteresis | Hold never exceeds one step | All cases in U-10 and U-11 | Max 1 step from natural value | n/a | Held in 10,337 voltage and 2,639 current cases, always exactly 1 step | PASS | Unit test
E-01 | Edge | Channel independence | One channel changes every cycle, the other fixed, both ways | Fixed channel never changes | n/a | 122 cycles: 0 changes on the fixed channel, 0 value mismatches | PASS | Simulation
E-02 | Edge | Large load steps | 0 -> 3.37 A -> 0 -> 4.89 A, 5 -> 12 -> 19.66 V | Shown in the same cycle | n/a | Shown immediately (hysteresis only acts on 1-step changes) | PASS | Simulation
E-03 | Edge | Noise at voltage full scale | ADC 1022.5 ±1 count, 100 cycles | Documented behavior | n/a | Alternates OL / 20.100-20.110V (50 changes). Outside the 0.8-20 V rating (F-04) | OBSERVATION | Simulation
E-04 | Edge | Noise at the 5.00 A limit | ADC 911.12 ±1 count, 100 cycles | Documented behavior | n/a | Alternates 5.000A / OL (48 changes). Outside the 4 A rating (F-04) | OBSERVATION | Simulation
E-05 | Edge | LED at the 0.00/0.01 A boundary | 6 mV ±1 count noise, 100 cycles | LED matches display | n/a | 0 inconsistencies, 0 display changes | PASS | Simulation
E-06 | Edge | Power-on with the alert already active | Temperature 0.689 V at boot | Alert on first cycle, LEDs off, clean return | n/a | Alert on first cycle, LEDs off, clean rows after cooling | PASS | Simulation
E-07 | Edge | Alert during over range | OL on CH1, then alert, then back | OL shown again after alert | n/a | Correct | PASS | Simulation
E-08 | Edge | Watchdog register state | Read WDTCSR during splash and after setup | Off during splash, 2 s after setup | n/a | During splash 0x00, after setup 0x0F | PASS | Simulation
E-09 | Edge | Irregular loop() calls | loop() every 1 ms and every 9 ms for 20 s | No measurement before 560 ms | n/a | Every 1 ms: all intervals 560 ms. Every 9 ms: all 567 ms (next call after 560) | PASS | Simulation
E-10 | Edge | Random soak | 3,000 cycles, random inputs, noise, dropouts, random temperature | Valid rows, correct LEDs and alert state | n/a | 347 alert cycles; 0 format, LED, alert-state or read-count errors | PASS | Simulation
E-11 | Edge | Worst-case sorting time | Samples in reverse order (worst case for insertion sort) | Well inside 560 ms and 2 s watchdog | n/a | +3.5 ms per 64-sample channel; worst-case cycle about 57 ms (see Memory & Timing) | PASS | Simulation
G-01 | Debug output | Serial Monitor format | DEBUG_FLAG 1: normal, OL and alert cycles | One temperature line then one line per channel | n/a | Correct; OL printed as "OL"; only temperature line during alert; 115200 baud | PASS | Simulation
V-11 | Voltage | Displayed range | Fine sweep ADC 0-1023 (0.02 steps) | 0.790 V up to the top of the ADC range, then OL | 0.790-20.130 V, never OL | 0.790-20.110 V; OL from ADC 1022.5 (20.160 V measured) | PASS | Simulation + model
V-12 | Voltage | Bench check against a Fluke 15B+ | 10 points 0.8-20 V on the real supply, final firmware | Every difference inside the Fluke tolerance | n/a | All 10 inside. Below 4 V (Fluke 1 mV range): -7 to -2 mV. Above 4 V the Fluke is not precise enough to decide (F-13) | PASS | Bench (hardware)

## Unit Tests
Unit tests: each firmware function called directly on the simulated chip
Every result compared with an independent float32 model of the intended behavior. Run on the release file.
Function | What was tested | Cases | Mismatches | Result
readAdcAverage() | Scripted sample sets: equal values, sort order, dropouts, spikes, limit 8/9 counts, two levels, random noise, 16 samples, 100 -> 64 cap, 1 sample | 15 | 0 | PASS
computeVoltage() | Every 10 mV step boundary (steps 79-2017) x offsets -3.1..+5 mV x previous state (none, -2, -1, same, +1), plus OL and 0.790 V threshold edges | 96962 | 0 | PASS
computeCurrent() | Every table midpoint and entry x offsets ±1.6 mV x previous state, plus 4451 mV / ADC 911-912 OL edges | 25001 | 0 | PASS
findNearestCurrentStep() | 0 to 4,451 mV in 0.25 mV steps against a plain nearest-value search | 17805 | 0 | PASS
formatChannelRow() | Every reachable voltage (0.000, 0.790-20.110 V) and current value, and all OL combinations | 2438 | 0 | PASS
computeVoltage() fine sweep | ADC 0 to 1023 in 0.02 count steps, fresh state each point, against the float32 calibration-line model | 51151 | 0 | PASS
computeVoltage() ramps | ADC 0 to 1023 and back in 0.05 count steps with the hysteresis state carried over: matches model, 0 reversals, never more than 1 step | 40922 | 0 | PASS
shouldVRefTemperatureAlertBeActive() | Both states x 8 voltages around 0.670 and 0.680 V | 16 | 0 | PASS
isAdjacentStep() | NO_STEP, equal, ±1, ±2 and extreme values | 12 | 0 | PASS
Total |  | 234322 | 0 | PASS
readAdcAverage() with 0 samples would return NaN; the static_assert "At least one sample" prevents it at compile time. Unreachable previous states (current step 501) are excluded.

## ADC Parity
Full ADC parity: v0.3.0 vs v1.0.0, every code 0-1023
Channel 1 row shown; channel 2 was identical in every step. Inputs are exact (noise-free) ADC codes. Voltage differences come from the calibration line that replaces the 7 offset bands.
ADC counts per volt |  | 204.7 | Op-amp gain |  | 4.036 | Constants from the firmware (ADC_COUNTS_PER_VOLT, OP_AMP_GAIN). v1.0.0 then applies displayed = measured x 0.99578 + 31 mV.
ADC code | ADC pin (mV) | Channel voltage (mV) | v0.3.0 voltage (V) | v1.0.0 voltage (V) | Voltage result | Current sense (mV) | v0.3.0 current (A) | v1.0.0 current (A) | Current result | v0.3.0 LED | v1.0.0 LED | LED result | v0.3.0 LCD row (current sweep) | v1.0.0 LCD row (current sweep)
0 | 0 | 0 | 0 | 0 | Identical | 0 | 0 | 0 | Identical | Off | Off | Same | ¦12.030V 0.000A ¦ | ¦12.030V 0.000A ¦
1 | 4.8852 | 19.7167 | 0 | 0 | Identical | 4.8852 | 0 | 0 | Identical | Off | Off | Same | ¦12.030V 0.000A ¦ | ¦12.030V 0.000A ¦
2 | 9.7704 | 39.4333 | 0 | 0 | Identical | 9.7704 | 0.01 | 0.01 | Identical | On | On | Same | ¦12.030V 0.010A ¦ | ¦12.030V 0.010A ¦
3 | 14.6556 | 59.15 | 0 | 0 | Identical | 14.6556 | 0.01 | 0.01 | Identical | On | On | Same | ¦12.030V 0.010A ¦ | ¦12.030V 0.010A ¦
4 | 19.5408 | 78.8666 | 0 | 0 | Identical | 19.5408 | 0.02 | 0.02 | Identical | On | On | Same | ¦12.030V 0.020A ¦ | ¦12.030V 0.020A ¦
5 | 24.426 | 98.5833 | 0 | 0 | Identical | 24.426 | 0.02 | 0.02 | Identical | On | On | Same | ¦12.030V 0.020A ¦ | ¦12.030V 0.020A ¦
6 | 29.3112 | 118.3 | 0 | 0 | Identical | 29.3112 | 0.03 | 0.03 | Identical | On | On | Same | ¦12.030V 0.030A ¦ | ¦12.030V 0.030A ¦
7 | 34.1964 | 138.0166 | 0 | 0 | Identical | 34.1964 | 0.03 | 0.03 | Identical | On | On | Same | ¦12.030V 0.030A ¦ | ¦12.030V 0.030A ¦
8 | 39.0816 | 157.7333 | 0 | 0 | Identical | 39.0816 | 0.04 | 0.04 | Identical | On | On | Same | ¦12.030V 0.040A ¦ | ¦12.030V 0.040A ¦
9 | 43.9668 | 177.4499 | 0 | 0 | Identical | 43.9668 | 0.04 | 0.05 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 0.040A ¦ | ¦12.030V 0.050A ¦
10 | 48.852 | 197.1666 | 0 | 0 | Identical | 48.852 | 0.05 | 0.05 | Identical | On | On | Same | ¦12.030V 0.050A ¦ | ¦12.030V 0.050A ¦
11 | 53.7372 | 216.8832 | 0 | 0 | Identical | 53.7372 | 0.06 | 0.06 | Identical | On | On | Same | ¦12.030V 0.060A ¦ | ¦12.030V 0.060A ¦
12 | 58.6224 | 236.5999 | 0 | 0 | Identical | 58.6224 | 0.06 | 0.06 | Identical | On | On | Same | ¦12.030V 0.060A ¦ | ¦12.030V 0.060A ¦
13 | 63.5076 | 256.3166 | 0 | 0 | Identical | 63.5076 | 0.07 | 0.07 | Identical | On | On | Same | ¦12.030V 0.070A ¦ | ¦12.030V 0.070A ¦
14 | 68.3928 | 276.0332 | 0 | 0 | Identical | 68.3928 | 0.07 | 0.07 | Identical | On | On | Same | ¦12.030V 0.070A ¦ | ¦12.030V 0.070A ¦
15 | 73.278 | 295.7499 | 0 | 0 | Identical | 73.278 | 0.08 | 0.08 | Identical | On | On | Same | ¦12.030V 0.080A ¦ | ¦12.030V 0.080A ¦
16 | 78.1632 | 315.4665 | 0 | 0 | Identical | 78.1632 | 0.08 | 0.08 | Identical | On | On | Same | ¦12.030V 0.080A ¦ | ¦12.030V 0.080A ¦
17 | 83.0484 | 335.1832 | 0 | 0 | Identical | 83.0484 | 0.09 | 0.09 | Identical | On | On | Same | ¦12.030V 0.090A ¦ | ¦12.030V 0.090A ¦
18 | 87.9336 | 354.8999 | 0 | 0 | Identical | 87.9336 | 0.09 | 0.09 | Identical | On | On | Same | ¦12.030V 0.090A ¦ | ¦12.030V 0.090A ¦
19 | 92.8188 | 374.6165 | 0 | 0 | Identical | 92.8188 | 0.1 | 0.1 | Identical | On | On | Same | ¦12.030V 0.100A ¦ | ¦12.030V 0.100A ¦
20 | 97.704 | 394.3332 | 0 | 0 | Identical | 97.704 | 0.1 | 0.11 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 0.100A ¦ | ¦12.030V 0.110A ¦
21 | 102.5892 | 414.0498 | 0 | 0 | Identical | 102.5892 | 0.11 | 0.11 | Identical | On | On | Same | ¦12.030V 0.110A ¦ | ¦12.030V 0.110A ¦
22 | 107.4744 | 433.7665 | 0 | 0 | Identical | 107.4744 | 0.12 | 0.12 | Identical | On | On | Same | ¦12.030V 0.120A ¦ | ¦12.030V 0.120A ¦
23 | 112.3596 | 453.4831 | 0 | 0 | Identical | 112.3596 | 0.12 | 0.12 | Identical | On | On | Same | ¦12.030V 0.120A ¦ | ¦12.030V 0.120A ¦
24 | 117.2447 | 473.1998 | 0 | 0 | Identical | 117.2447 | 0.13 | 0.13 | Identical | On | On | Same | ¦12.030V 0.130A ¦ | ¦12.030V 0.130A ¦
25 | 122.1299 | 492.9165 | 0 | 0 | Identical | 122.1299 | 0.13 | 0.13 | Identical | On | On | Same | ¦12.030V 0.130A ¦ | ¦12.030V 0.130A ¦
26 | 127.0151 | 512.6331 | 0 | 0 | Identical | 127.0151 | 0.14 | 0.14 | Identical | On | On | Same | ¦12.030V 0.140A ¦ | ¦12.030V 0.140A ¦
27 | 131.9003 | 532.3498 | 0 | 0 | Identical | 131.9003 | 0.14 | 0.14 | Identical | On | On | Same | ¦12.030V 0.140A ¦ | ¦12.030V 0.140A ¦
28 | 136.7855 | 552.0664 | 0 | 0 | Identical | 136.7855 | 0.15 | 0.15 | Identical | On | On | Same | ¦12.030V 0.150A ¦ | ¦12.030V 0.150A ¦
29 | 141.6707 | 571.7831 | 0 | 0 | Identical | 141.6707 | 0.15 | 0.16 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 0.150A ¦ | ¦12.030V 0.160A ¦
30 | 146.5559 | 591.4998 | 0 | 0 | Identical | 146.5559 | 0.16 | 0.16 | Identical | On | On | Same | ¦12.030V 0.160A ¦ | ¦12.030V 0.160A ¦
31 | 151.4411 | 611.2164 | 0 | 0 | Identical | 151.4411 | 0.17 | 0.17 | Identical | On | On | Same | ¦12.030V 0.170A ¦ | ¦12.030V 0.170A ¦
32 | 156.3263 | 630.9331 | 0 | 0 | Identical | 156.3263 | 0.17 | 0.17 | Identical | On | On | Same | ¦12.030V 0.170A ¦ | ¦12.030V 0.170A ¦
33 | 161.2115 | 650.6497 | 0 | 0 | Identical | 161.2115 | 0.18 | 0.18 | Identical | On | On | Same | ¦12.030V 0.180A ¦ | ¦12.030V 0.180A ¦
34 | 166.0967 | 670.3664 | 0 | 0 | Identical | 166.0967 | 0.18 | 0.18 | Identical | On | On | Same | ¦12.030V 0.180A ¦ | ¦12.030V 0.180A ¦
35 | 170.9819 | 690.083 | 0 | 0 | Identical | 170.9819 | 0.19 | 0.19 | Identical | On | On | Same | ¦12.030V 0.190A ¦ | ¦12.030V 0.190A ¦
36 | 175.8671 | 709.7997 | 0 | 0 | Identical | 175.8671 | 0.19 | 0.19 | Identical | On | On | Same | ¦12.030V 0.190A ¦ | ¦12.030V 0.190A ¦
37 | 180.7523 | 729.5164 | 0 | 0 | Identical | 180.7523 | 0.2 | 0.2 | Identical | On | On | Same | ¦12.030V 0.200A ¦ | ¦12.030V 0.200A ¦
38 | 185.6375 | 749.233 | 0 | 0 | Identical | 185.6375 | 0.2 | 0.2 | Identical | On | On | Same | ¦12.030V 0.200A ¦ | ¦12.030V 0.200A ¦
39 | 190.5227 | 768.9497 | 0.79 | 0.8 | Within 10 mV (calibration line) | 190.5227 | 0.21 | 0.21 | Identical | On | On | Same | ¦12.030V 0.210A ¦ | ¦12.030V 0.210A ¦
40 | 195.4079 | 788.6663 | 0.79 | 0.82 | Low range fixed (F-02) | 195.4079 | 0.21 | 0.21 | Identical | On | On | Same | ¦12.030V 0.210A ¦ | ¦12.030V 0.210A ¦
41 | 200.2931 | 808.383 | 0.8 | 0.84 | Low range fixed (F-02) | 200.2931 | 0.22 | 0.22 | Identical | On | On | Same | ¦12.030V 0.220A ¦ | ¦12.030V 0.220A ¦
42 | 205.1783 | 828.0997 | 0.85 | 0.86 | Within 10 mV (calibration line) | 205.1783 | 0.22 | 0.23 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 0.220A ¦ | ¦12.030V 0.230A ¦
43 | 210.0635 | 847.8163 | 0.87 | 0.88 | Within 10 mV (calibration line) | 210.0635 | 0.23 | 0.23 | Identical | On | On | Same | ¦12.030V 0.230A ¦ | ¦12.030V 0.230A ¦
44 | 214.9487 | 867.533 | 0.89 | 0.89 | Identical | 214.9487 | 0.24 | 0.24 | Identical | On | On | Same | ¦12.030V 0.240A ¦ | ¦12.030V 0.240A ¦
45 | 219.8339 | 887.2496 | 0.91 | 0.91 | Identical | 219.8339 | 0.24 | 0.24 | Identical | On | On | Same | ¦12.030V 0.240A ¦ | ¦12.030V 0.240A ¦
46 | 224.7191 | 906.9663 | 0.93 | 0.93 | Identical | 224.7191 | 0.25 | 0.25 | Identical | On | On | Same | ¦12.030V 0.250A ¦ | ¦12.030V 0.250A ¦
47 | 229.6043 | 926.683 | 0.95 | 0.95 | Identical | 229.6043 | 0.25 | 0.25 | Identical | On | On | Same | ¦12.030V 0.250A ¦ | ¦12.030V 0.250A ¦
48 | 234.4895 | 946.3996 | 0.97 | 0.97 | Identical | 234.4895 | 0.26 | 0.26 | Identical | On | On | Same | ¦12.030V 0.260A ¦ | ¦12.030V 0.260A ¦
49 | 239.3747 | 966.1163 | 0.99 | 0.99 | Identical | 239.3747 | 0.26 | 0.26 | Identical | On | On | Same | ¦12.030V 0.260A ¦ | ¦12.030V 0.260A ¦
50 | 244.2599 | 985.8329 | 1.01 | 1.01 | Identical | 244.2599 | 0.27 | 0.27 | Identical | On | On | Same | ¦12.030V 0.270A ¦ | ¦12.030V 0.270A ¦
51 | 249.1451 | 1005.5496 | 1.03 | 1.03 | Identical | 249.1451 | 0.27 | 0.27 | Identical | On | On | Same | ¦12.030V 0.270A ¦ | ¦12.030V 0.270A ¦
52 | 254.0303 | 1025.2662 | 1.05 | 1.05 | Identical | 254.0303 | 0.28 | 0.28 | Identical | On | On | Same | ¦12.030V 0.280A ¦ | ¦12.030V 0.280A ¦
53 | 258.9155 | 1044.9829 | 1.07 | 1.07 | Identical | 258.9155 | 0.28 | 0.29 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 0.280A ¦ | ¦12.030V 0.290A ¦
54 | 263.8007 | 1064.6996 | 1.09 | 1.09 | Identical | 263.8007 | 0.29 | 0.29 | Identical | On | On | Same | ¦12.030V 0.290A ¦ | ¦12.030V 0.290A ¦
55 | 268.6859 | 1084.4162 | 1.11 | 1.11 | Identical | 268.6859 | 0.3 | 0.3 | Identical | On | On | Same | ¦12.030V 0.300A ¦ | ¦12.030V 0.300A ¦
56 | 273.5711 | 1104.1329 | 1.13 | 1.13 | Identical | 273.5711 | 0.3 | 0.3 | Identical | On | On | Same | ¦12.030V 0.300A ¦ | ¦12.030V 0.300A ¦
57 | 278.4563 | 1123.8495 | 1.15 | 1.15 | Identical | 278.4563 | 0.31 | 0.31 | Identical | On | On | Same | ¦12.030V 0.310A ¦ | ¦12.030V 0.310A ¦
58 | 283.3415 | 1143.5662 | 1.17 | 1.17 | Identical | 283.3415 | 0.31 | 0.31 | Identical | On | On | Same | ¦12.030V 0.310A ¦ | ¦12.030V 0.310A ¦
59 | 288.2267 | 1163.2829 | 1.19 | 1.19 | Identical | 288.2267 | 0.32 | 0.32 | Identical | On | On | Same | ¦12.030V 0.320A ¦ | ¦12.030V 0.320A ¦
60 | 293.1119 | 1182.9995 | 1.21 | 1.21 | Identical | 293.1119 | 0.32 | 0.32 | Identical | On | On | Same | ¦12.030V 0.320A ¦ | ¦12.030V 0.320A ¦
61 | 297.9971 | 1202.7162 | 1.23 | 1.23 | Identical | 297.9971 | 0.33 | 0.33 | Identical | On | On | Same | ¦12.030V 0.330A ¦ | ¦12.030V 0.330A ¦
62 | 302.8823 | 1222.4328 | 1.25 | 1.25 | Identical | 302.8823 | 0.33 | 0.34 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 0.330A ¦ | ¦12.030V 0.340A ¦
63 | 307.7675 | 1242.1495 | 1.27 | 1.27 | Identical | 307.7675 | 0.34 | 0.34 | Identical | On | On | Same | ¦12.030V 0.340A ¦ | ¦12.030V 0.340A ¦
64 | 312.6527 | 1261.8661 | 1.29 | 1.29 | Identical | 312.6527 | 0.35 | 0.35 | Identical | On | On | Same | ¦12.030V 0.350A ¦ | ¦12.030V 0.350A ¦
65 | 317.5379 | 1281.5828 | 1.31 | 1.31 | Identical | 317.5379 | 0.35 | 0.35 | Identical | On | On | Same | ¦12.030V 0.350A ¦ | ¦12.030V 0.350A ¦
66 | 322.4231 | 1301.2995 | 1.33 | 1.33 | Identical | 322.4231 | 0.36 | 0.36 | Identical | On | On | Same | ¦12.030V 0.360A ¦ | ¦12.030V 0.360A ¦
67 | 327.3083 | 1321.0161 | 1.35 | 1.35 | Identical | 327.3083 | 0.36 | 0.36 | Identical | On | On | Same | ¦12.030V 0.360A ¦ | ¦12.030V 0.360A ¦
68 | 332.1935 | 1340.7328 | 1.37 | 1.37 | Identical | 332.1935 | 0.37 | 0.37 | Identical | On | On | Same | ¦12.030V 0.370A ¦ | ¦12.030V 0.370A ¦
69 | 337.0787 | 1360.4494 | 1.39 | 1.39 | Identical | 337.0787 | 0.37 | 0.37 | Identical | On | On | Same | ¦12.030V 0.370A ¦ | ¦12.030V 0.370A ¦
70 | 341.9638 | 1380.1661 | 1.41 | 1.41 | Identical | 341.9638 | 0.38 | 0.38 | Identical | On | On | Same | ¦12.030V 0.380A ¦ | ¦12.030V 0.380A ¦
71 | 346.849 | 1399.8828 | 1.42 | 1.42 | Identical | 346.849 | 0.38 | 0.38 | Identical | On | On | Same | ¦12.030V 0.380A ¦ | ¦12.030V 0.380A ¦
72 | 351.7342 | 1419.5994 | 1.44 | 1.44 | Identical | 351.7342 | 0.39 | 0.39 | Identical | On | On | Same | ¦12.030V 0.390A ¦ | ¦12.030V 0.390A ¦
73 | 356.6194 | 1439.3161 | 1.46 | 1.46 | Identical | 356.6194 | 0.39 | 0.4 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 0.390A ¦ | ¦12.030V 0.400A ¦
74 | 361.5046 | 1459.0327 | 1.48 | 1.48 | Identical | 361.5046 | 0.4 | 0.4 | Identical | On | On | Same | ¦12.030V 0.400A ¦ | ¦12.030V 0.400A ¦
75 | 366.3898 | 1478.7494 | 1.5 | 1.5 | Identical | 366.3898 | 0.41 | 0.41 | Identical | On | On | Same | ¦12.030V 0.410A ¦ | ¦12.030V 0.410A ¦
76 | 371.275 | 1498.466 | 1.52 | 1.52 | Identical | 371.275 | 0.41 | 0.41 | Identical | On | On | Same | ¦12.030V 0.410A ¦ | ¦12.030V 0.410A ¦
77 | 376.1602 | 1518.1827 | 1.54 | 1.54 | Identical | 376.1602 | 0.42 | 0.42 | Identical | On | On | Same | ¦12.030V 0.420A ¦ | ¦12.030V 0.420A ¦
78 | 381.0454 | 1537.8994 | 1.56 | 1.56 | Identical | 381.0454 | 0.42 | 0.42 | Identical | On | On | Same | ¦12.030V 0.420A ¦ | ¦12.030V 0.420A ¦
79 | 385.9306 | 1557.616 | 1.58 | 1.58 | Identical | 385.9306 | 0.43 | 0.43 | Identical | On | On | Same | ¦12.030V 0.430A ¦ | ¦12.030V 0.430A ¦
80 | 390.8158 | 1577.3327 | 1.6 | 1.6 | Identical | 390.8158 | 0.43 | 0.43 | Identical | On | On | Same | ¦12.030V 0.430A ¦ | ¦12.030V 0.430A ¦
81 | 395.701 | 1597.0493 | 1.62 | 1.62 | Identical | 395.701 | 0.44 | 0.44 | Identical | On | On | Same | ¦12.030V 0.440A ¦ | ¦12.030V 0.440A ¦
82 | 400.5862 | 1616.766 | 1.64 | 1.64 | Identical | 400.5862 | 0.44 | 0.44 | Identical | On | On | Same | ¦12.030V 0.440A ¦ | ¦12.030V 0.440A ¦
83 | 405.4714 | 1636.4827 | 1.66 | 1.66 | Identical | 405.4714 | 0.45 | 0.45 | Identical | On | On | Same | ¦12.030V 0.450A ¦ | ¦12.030V 0.450A ¦
84 | 410.3566 | 1656.1993 | 1.68 | 1.68 | Identical | 410.3566 | 0.45 | 0.45 | Identical | On | On | Same | ¦12.030V 0.450A ¦ | ¦12.030V 0.450A ¦
85 | 415.2418 | 1675.916 | 1.7 | 1.7 | Identical | 415.2418 | 0.46 | 0.46 | Identical | On | On | Same | ¦12.030V 0.460A ¦ | ¦12.030V 0.460A ¦
86 | 420.127 | 1695.6326 | 1.72 | 1.72 | Identical | 420.127 | 0.47 | 0.47 | Identical | On | On | Same | ¦12.030V 0.470A ¦ | ¦12.030V 0.470A ¦
87 | 425.0122 | 1715.3493 | 1.74 | 1.74 | Identical | 425.0122 | 0.47 | 0.47 | Identical | On | On | Same | ¦12.030V 0.470A ¦ | ¦12.030V 0.470A ¦
88 | 429.8974 | 1735.066 | 1.76 | 1.76 | Identical | 429.8974 | 0.48 | 0.48 | Identical | On | On | Same | ¦12.030V 0.480A ¦ | ¦12.030V 0.480A ¦
89 | 434.7826 | 1754.7826 | 1.78 | 1.78 | Identical | 434.7826 | 0.48 | 0.48 | Identical | On | On | Same | ¦12.030V 0.480A ¦ | ¦12.030V 0.480A ¦
90 | 439.6678 | 1774.4993 | 1.8 | 1.8 | Identical | 439.6678 | 0.49 | 0.49 | Identical | On | On | Same | ¦12.030V 0.490A ¦ | ¦12.030V 0.490A ¦
91 | 444.553 | 1794.2159 | 1.82 | 1.82 | Identical | 444.553 | 0.49 | 0.49 | Identical | On | On | Same | ¦12.030V 0.490A ¦ | ¦12.030V 0.490A ¦
92 | 449.4382 | 1813.9326 | 1.84 | 1.84 | Identical | 449.4382 | 0.5 | 0.5 | Identical | On | On | Same | ¦12.030V 0.500A ¦ | ¦12.030V 0.500A ¦
93 | 454.3234 | 1833.6492 | 1.86 | 1.86 | Identical | 454.3234 | 0.5 | 0.5 | Identical | On | On | Same | ¦12.030V 0.500A ¦ | ¦12.030V 0.500A ¦
94 | 459.2086 | 1853.3659 | 1.88 | 1.88 | Identical | 459.2086 | 0.51 | 0.51 | Identical | On | On | Same | ¦12.030V 0.510A ¦ | ¦12.030V 0.510A ¦
95 | 464.0938 | 1873.0826 | 1.9 | 1.9 | Identical | 464.0938 | 0.52 | 0.52 | Identical | On | On | Same | ¦12.030V 0.520A ¦ | ¦12.030V 0.520A ¦
96 | 468.979 | 1892.7992 | 1.92 | 1.92 | Identical | 468.979 | 0.52 | 0.52 | Identical | On | On | Same | ¦12.030V 0.520A ¦ | ¦12.030V 0.520A ¦
97 | 473.8642 | 1912.5159 | 1.94 | 1.94 | Identical | 473.8642 | 0.53 | 0.53 | Identical | On | On | Same | ¦12.030V 0.530A ¦ | ¦12.030V 0.530A ¦
98 | 478.7494 | 1932.2325 | 1.96 | 1.96 | Identical | 478.7494 | 0.53 | 0.53 | Identical | On | On | Same | ¦12.030V 0.530A ¦ | ¦12.030V 0.530A ¦
99 | 483.6346 | 1951.9492 | 1.98 | 1.97 | Within 10 mV (calibration line) | 483.6346 | 0.54 | 0.54 | Identical | On | On | Same | ¦12.030V 0.540A ¦ | ¦12.030V 0.540A ¦
100 | 488.5198 | 1971.6659 | 2 | 1.99 | Within 10 mV (calibration line) | 488.5198 | 0.54 | 0.54 | Identical | On | On | Same | ¦12.030V 0.540A ¦ | ¦12.030V 0.540A ¦
101 | 493.405 | 1991.3825 | 2.02 | 2.01 | Within 10 mV (calibration line) | 493.405 | 0.55 | 0.55 | Identical | On | On | Same | ¦12.030V 0.550A ¦ | ¦12.030V 0.550A ¦
102 | 498.2902 | 2011.0992 | 2.04 | 2.03 | Within 10 mV (calibration line) | 498.2902 | 0.55 | 0.55 | Identical | On | On | Same | ¦12.030V 0.550A ¦ | ¦12.030V 0.550A ¦
103 | 503.1754 | 2030.8158 | 2.06 | 2.05 | Within 10 mV (calibration line) | 503.1754 | 0.56 | 0.56 | Identical | On | On | Same | ¦12.030V 0.560A ¦ | ¦12.030V 0.560A ¦
104 | 508.0606 | 2050.5325 | 2.08 | 2.07 | Within 10 mV (calibration line) | 508.0606 | 0.56 | 0.56 | Identical | On | On | Same | ¦12.030V 0.560A ¦ | ¦12.030V 0.560A ¦
105 | 512.9458 | 2070.2491 | 2.1 | 2.09 | Within 10 mV (calibration line) | 512.9458 | 0.57 | 0.57 | Identical | On | On | Same | ¦12.030V 0.570A ¦ | ¦12.030V 0.570A ¦
106 | 517.831 | 2089.9658 | 2.11 | 2.11 | Identical | 517.831 | 0.57 | 0.58 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 0.570A ¦ | ¦12.030V 0.580A ¦
107 | 522.7162 | 2109.6825 | 2.13 | 2.13 | Identical | 522.7162 | 0.58 | 0.58 | Identical | On | On | Same | ¦12.030V 0.580A ¦ | ¦12.030V 0.580A ¦
108 | 527.6014 | 2129.3991 | 2.15 | 2.15 | Identical | 527.6014 | 0.59 | 0.59 | Identical | On | On | Same | ¦12.030V 0.590A ¦ | ¦12.030V 0.590A ¦
109 | 532.4866 | 2149.1158 | 2.17 | 2.17 | Identical | 532.4866 | 0.59 | 0.59 | Identical | On | On | Same | ¦12.030V 0.590A ¦ | ¦12.030V 0.590A ¦
110 | 537.3718 | 2168.8324 | 2.19 | 2.19 | Identical | 537.3718 | 0.6 | 0.6 | Identical | On | On | Same | ¦12.030V 0.600A ¦ | ¦12.030V 0.600A ¦
111 | 542.257 | 2188.5491 | 2.21 | 2.21 | Identical | 542.257 | 0.6 | 0.6 | Identical | On | On | Same | ¦12.030V 0.600A ¦ | ¦12.030V 0.600A ¦
112 | 547.1422 | 2208.2658 | 2.23 | 2.23 | Identical | 547.1422 | 0.61 | 0.61 | Identical | On | On | Same | ¦12.030V 0.610A ¦ | ¦12.030V 0.610A ¦
113 | 552.0274 | 2227.9824 | 2.25 | 2.25 | Identical | 552.0274 | 0.61 | 0.61 | Identical | On | On | Same | ¦12.030V 0.610A ¦ | ¦12.030V 0.610A ¦
114 | 556.9126 | 2247.6991 | 2.27 | 2.27 | Identical | 556.9126 | 0.62 | 0.62 | Identical | On | On | Same | ¦12.030V 0.620A ¦ | ¦12.030V 0.620A ¦
115 | 561.7978 | 2267.4157 | 2.29 | 2.29 | Identical | 561.7978 | 0.62 | 0.62 | Identical | On | On | Same | ¦12.030V 0.620A ¦ | ¦12.030V 0.620A ¦
116 | 566.683 | 2287.1324 | 2.31 | 2.31 | Identical | 566.683 | 0.63 | 0.63 | Identical | On | On | Same | ¦12.030V 0.630A ¦ | ¦12.030V 0.630A ¦
117 | 571.5681 | 2306.849 | 2.33 | 2.33 | Identical | 571.5681 | 0.63 | 0.64 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 0.630A ¦ | ¦12.030V 0.640A ¦
118 | 576.4533 | 2326.5657 | 2.35 | 2.35 | Identical | 576.4533 | 0.64 | 0.64 | Identical | On | On | Same | ¦12.030V 0.640A ¦ | ¦12.030V 0.640A ¦
119 | 581.3385 | 2346.2824 | 2.37 | 2.37 | Identical | 581.3385 | 0.65 | 0.65 | Identical | On | On | Same | ¦12.030V 0.650A ¦ | ¦12.030V 0.650A ¦
120 | 586.2237 | 2365.999 | 2.39 | 2.39 | Identical | 586.2237 | 0.65 | 0.65 | Identical | On | On | Same | ¦12.030V 0.650A ¦ | ¦12.030V 0.650A ¦
121 | 591.1089 | 2385.7157 | 2.41 | 2.41 | Identical | 591.1089 | 0.66 | 0.66 | Identical | On | On | Same | ¦12.030V 0.660A ¦ | ¦12.030V 0.660A ¦
122 | 595.9941 | 2405.4323 | 2.43 | 2.43 | Identical | 595.9941 | 0.66 | 0.66 | Identical | On | On | Same | ¦12.030V 0.660A ¦ | ¦12.030V 0.660A ¦
123 | 600.8793 | 2425.149 | 2.45 | 2.45 | Identical | 600.8793 | 0.67 | 0.67 | Identical | On | On | Same | ¦12.030V 0.670A ¦ | ¦12.030V 0.670A ¦
124 | 605.7645 | 2444.8657 | 2.47 | 2.47 | Identical | 605.7645 | 0.67 | 0.67 | Identical | On | On | Same | ¦12.030V 0.670A ¦ | ¦12.030V 0.670A ¦
125 | 610.6497 | 2464.5823 | 2.49 | 2.49 | Identical | 610.6497 | 0.68 | 0.68 | Identical | On | On | Same | ¦12.030V 0.680A ¦ | ¦12.030V 0.680A ¦
126 | 615.5349 | 2484.299 | 2.51 | 2.5 | Within 10 mV (calibration line) | 615.5349 | 0.68 | 0.69 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 0.680A ¦ | ¦12.030V 0.690A ¦
127 | 620.4201 | 2504.0156 | 2.53 | 2.52 | Within 10 mV (calibration line) | 620.4201 | 0.69 | 0.69 | Identical | On | On | Same | ¦12.030V 0.690A ¦ | ¦12.030V 0.690A ¦
128 | 625.3053 | 2523.7323 | 2.55 | 2.54 | Within 10 mV (calibration line) | 625.3053 | 0.7 | 0.7 | Identical | On | On | Same | ¦12.030V 0.700A ¦ | ¦12.030V 0.700A ¦
129 | 630.1905 | 2543.4489 | 2.57 | 2.56 | Within 10 mV (calibration line) | 630.1905 | 0.71 | 0.71 | Identical | On | On | Same | ¦12.030V 0.710A ¦ | ¦12.030V 0.710A ¦
130 | 635.0757 | 2563.1656 | 2.59 | 2.58 | Within 10 mV (calibration line) | 635.0757 | 0.71 | 0.71 | Identical | On | On | Same | ¦12.030V 0.710A ¦ | ¦12.030V 0.710A ¦
131 | 639.9609 | 2582.8823 | 2.61 | 2.6 | Within 10 mV (calibration line) | 639.9609 | 0.72 | 0.72 | Identical | On | On | Same | ¦12.030V 0.720A ¦ | ¦12.030V 0.720A ¦
132 | 644.8461 | 2602.5989 | 2.63 | 2.62 | Within 10 mV (calibration line) | 644.8461 | 0.72 | 0.72 | Identical | On | On | Same | ¦12.030V 0.720A ¦ | ¦12.030V 0.720A ¦
133 | 649.7313 | 2622.3156 | 2.65 | 2.64 | Within 10 mV (calibration line) | 649.7313 | 0.73 | 0.73 | Identical | On | On | Same | ¦12.030V 0.730A ¦ | ¦12.030V 0.730A ¦
134 | 654.6165 | 2642.0322 | 2.67 | 2.66 | Within 10 mV (calibration line) | 654.6165 | 0.73 | 0.73 | Identical | On | On | Same | ¦12.030V 0.730A ¦ | ¦12.030V 0.730A ¦
135 | 659.5017 | 2661.7489 | 2.69 | 2.68 | Within 10 mV (calibration line) | 659.5017 | 0.74 | 0.74 | Identical | On | On | Same | ¦12.030V 0.740A ¦ | ¦12.030V 0.740A ¦
136 | 664.3869 | 2681.4656 | 2.71 | 2.7 | Within 10 mV (calibration line) | 664.3869 | 0.74 | 0.74 | Identical | On | On | Same | ¦12.030V 0.740A ¦ | ¦12.030V 0.740A ¦
137 | 669.2721 | 2701.1822 | 2.73 | 2.72 | Within 10 mV (calibration line) | 669.2721 | 0.75 | 0.75 | Identical | On | On | Same | ¦12.030V 0.750A ¦ | ¦12.030V 0.750A ¦
138 | 674.1573 | 2720.8989 | 2.75 | 2.74 | Within 10 mV (calibration line) | 674.1573 | 0.75 | 0.75 | Identical | On | On | Same | ¦12.030V 0.750A ¦ | ¦12.030V 0.750A ¦
139 | 679.0425 | 2740.6155 | 2.77 | 2.76 | Within 10 mV (calibration line) | 679.0425 | 0.76 | 0.76 | Identical | On | On | Same | ¦12.030V 0.760A ¦ | ¦12.030V 0.760A ¦
140 | 683.9277 | 2760.3322 | 2.79 | 2.78 | Within 10 mV (calibration line) | 683.9277 | 0.76 | 0.77 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 0.760A ¦ | ¦12.030V 0.770A ¦
141 | 688.8129 | 2780.0489 | 2.81 | 2.8 | Within 10 mV (calibration line) | 688.8129 | 0.77 | 0.77 | Identical | On | On | Same | ¦12.030V 0.770A ¦ | ¦12.030V 0.770A ¦
142 | 693.6981 | 2799.7655 | 2.82 | 2.82 | Identical | 693.6981 | 0.78 | 0.78 | Identical | On | On | Same | ¦12.030V 0.780A ¦ | ¦12.030V 0.780A ¦
143 | 698.5833 | 2819.4822 | 2.84 | 2.84 | Identical | 698.5833 | 0.78 | 0.78 | Identical | On | On | Same | ¦12.030V 0.780A ¦ | ¦12.030V 0.780A ¦
144 | 703.4685 | 2839.1988 | 2.86 | 2.86 | Identical | 703.4685 | 0.79 | 0.79 | Identical | On | On | Same | ¦12.030V 0.790A ¦ | ¦12.030V 0.790A ¦
145 | 708.3537 | 2858.9155 | 2.88 | 2.88 | Identical | 708.3537 | 0.79 | 0.79 | Identical | On | On | Same | ¦12.030V 0.790A ¦ | ¦12.030V 0.790A ¦
146 | 713.2389 | 2878.6321 | 2.9 | 2.9 | Identical | 713.2389 | 0.8 | 0.8 | Identical | On | On | Same | ¦12.030V 0.800A ¦ | ¦12.030V 0.800A ¦
147 | 718.1241 | 2898.3488 | 2.92 | 2.92 | Identical | 718.1241 | 0.8 | 0.8 | Identical | On | On | Same | ¦12.030V 0.800A ¦ | ¦12.030V 0.800A ¦
148 | 723.0093 | 2918.0655 | 2.94 | 2.94 | Identical | 723.0093 | 0.81 | 0.81 | Identical | On | On | Same | ¦12.030V 0.810A ¦ | ¦12.030V 0.810A ¦
149 | 727.8945 | 2937.7821 | 2.96 | 2.96 | Identical | 727.8945 | 0.81 | 0.81 | Identical | On | On | Same | ¦12.030V 0.810A ¦ | ¦12.030V 0.810A ¦
150 | 732.7797 | 2957.4988 | 2.98 | 2.98 | Identical | 732.7797 | 0.82 | 0.82 | Identical | On | On | Same | ¦12.030V 0.820A ¦ | ¦12.030V 0.820A ¦
151 | 737.6649 | 2977.2154 | 3 | 3 | Identical | 737.6649 | 0.82 | 0.83 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 0.820A ¦ | ¦12.030V 0.830A ¦
152 | 742.5501 | 2996.9321 | 3.02 | 3.02 | Identical | 742.5501 | 0.83 | 0.83 | Identical | On | On | Same | ¦12.030V 0.830A ¦ | ¦12.030V 0.830A ¦
153 | 747.4353 | 3016.6488 | 3.04 | 3.03 | Within 10 mV (calibration line) | 747.4353 | 0.84 | 0.84 | Identical | On | On | Same | ¦12.030V 0.840A ¦ | ¦12.030V 0.840A ¦
154 | 752.3205 | 3036.3654 | 3.06 | 3.05 | Within 10 mV (calibration line) | 752.3205 | 0.84 | 0.84 | Identical | On | On | Same | ¦12.030V 0.840A ¦ | ¦12.030V 0.840A ¦
155 | 757.2057 | 3056.0821 | 3.08 | 3.07 | Within 10 mV (calibration line) | 757.2057 | 0.85 | 0.85 | Identical | On | On | Same | ¦12.030V 0.850A ¦ | ¦12.030V 0.850A ¦
156 | 762.0909 | 3075.7987 | 3.1 | 3.09 | Within 10 mV (calibration line) | 762.0909 | 0.85 | 0.85 | Identical | On | On | Same | ¦12.030V 0.850A ¦ | ¦12.030V 0.850A ¦
157 | 766.9761 | 3095.5154 | 3.12 | 3.11 | Within 10 mV (calibration line) | 766.9761 | 0.86 | 0.86 | Identical | On | On | Same | ¦12.030V 0.860A ¦ | ¦12.030V 0.860A ¦
158 | 771.8613 | 3115.232 | 3.14 | 3.13 | Within 10 mV (calibration line) | 771.8613 | 0.86 | 0.86 | Identical | On | On | Same | ¦12.030V 0.860A ¦ | ¦12.030V 0.860A ¦
159 | 776.7465 | 3134.9487 | 3.16 | 3.15 | Within 10 mV (calibration line) | 776.7465 | 0.87 | 0.87 | Identical | On | On | Same | ¦12.030V 0.870A ¦ | ¦12.030V 0.870A ¦
160 | 781.6317 | 3154.6654 | 3.18 | 3.17 | Within 10 mV (calibration line) | 781.6317 | 0.87 | 0.88 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 0.870A ¦ | ¦12.030V 0.880A ¦
161 | 786.5169 | 3174.382 | 3.2 | 3.19 | Within 10 mV (calibration line) | 786.5169 | 0.88 | 0.88 | Identical | On | On | Same | ¦12.030V 0.880A ¦ | ¦12.030V 0.880A ¦
162 | 791.4021 | 3194.0987 | 3.22 | 3.21 | Within 10 mV (calibration line) | 791.4021 | 0.89 | 0.89 | Identical | On | On | Same | ¦12.030V 0.890A ¦ | ¦12.030V 0.890A ¦
163 | 796.2872 | 3213.8153 | 3.24 | 3.23 | Within 10 mV (calibration line) | 796.2872 | 0.89 | 0.89 | Identical | On | On | Same | ¦12.030V 0.890A ¦ | ¦12.030V 0.890A ¦
164 | 801.1724 | 3233.532 | 3.26 | 3.25 | Within 10 mV (calibration line) | 801.1724 | 0.9 | 0.9 | Identical | On | On | Same | ¦12.030V 0.900A ¦ | ¦12.030V 0.900A ¦
165 | 806.0576 | 3253.2487 | 3.28 | 3.27 | Within 10 mV (calibration line) | 806.0576 | 0.9 | 0.9 | Identical | On | On | Same | ¦12.030V 0.900A ¦ | ¦12.030V 0.900A ¦
166 | 810.9428 | 3272.9653 | 3.3 | 3.29 | Within 10 mV (calibration line) | 810.9428 | 0.91 | 0.91 | Identical | On | On | Same | ¦12.030V 0.910A ¦ | ¦12.030V 0.910A ¦
167 | 815.828 | 3292.682 | 3.32 | 3.31 | Within 10 mV (calibration line) | 815.828 | 0.91 | 0.91 | Identical | On | On | Same | ¦12.030V 0.910A ¦ | ¦12.030V 0.910A ¦
168 | 820.7132 | 3312.3986 | 3.34 | 3.33 | Within 10 mV (calibration line) | 820.7132 | 0.92 | 0.92 | Identical | On | On | Same | ¦12.030V 0.920A ¦ | ¦12.030V 0.920A ¦
169 | 825.5984 | 3332.1153 | 3.36 | 3.35 | Within 10 mV (calibration line) | 825.5984 | 0.92 | 0.93 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 0.920A ¦ | ¦12.030V 0.930A ¦
170 | 830.4836 | 3351.8319 | 3.38 | 3.37 | Within 10 mV (calibration line) | 830.4836 | 0.93 | 0.93 | Identical | On | On | Same | ¦12.030V 0.930A ¦ | ¦12.030V 0.930A ¦
171 | 835.3688 | 3371.5486 | 3.4 | 3.39 | Within 10 mV (calibration line) | 835.3688 | 0.94 | 0.94 | Identical | On | On | Same | ¦12.030V 0.940A ¦ | ¦12.030V 0.940A ¦
172 | 840.254 | 3391.2653 | 3.42 | 3.41 | Within 10 mV (calibration line) | 840.254 | 0.94 | 0.94 | Identical | On | On | Same | ¦12.030V 0.940A ¦ | ¦12.030V 0.940A ¦
173 | 845.1392 | 3410.9819 | 3.44 | 3.43 | Within 10 mV (calibration line) | 845.1392 | 0.95 | 0.95 | Identical | On | On | Same | ¦12.030V 0.950A ¦ | ¦12.030V 0.950A ¦
174 | 850.0244 | 3430.6986 | 3.46 | 3.45 | Within 10 mV (calibration line) | 850.0244 | 0.95 | 0.95 | Identical | On | On | Same | ¦12.030V 0.950A ¦ | ¦12.030V 0.950A ¦
175 | 854.9096 | 3450.4152 | 3.48 | 3.47 | Within 10 mV (calibration line) | 854.9096 | 0.96 | 0.96 | Identical | On | On | Same | ¦12.030V 0.960A ¦ | ¦12.030V 0.960A ¦
176 | 859.7948 | 3470.1319 | 3.5 | 3.49 | Within 10 mV (calibration line) | 859.7948 | 0.96 | 0.96 | Identical | On | On | Same | ¦12.030V 0.960A ¦ | ¦12.030V 0.960A ¦
177 | 864.68 | 3489.8486 | 3.51 | 3.51 | Identical | 864.68 | 0.97 | 0.97 | Identical | On | On | Same | ¦12.030V 0.970A ¦ | ¦12.030V 0.970A ¦
178 | 869.5652 | 3509.5652 | 3.53 | 3.53 | Identical | 869.5652 | 0.97 | 0.97 | Identical | On | On | Same | ¦12.030V 0.970A ¦ | ¦12.030V 0.970A ¦
179 | 874.4504 | 3529.2819 | 3.55 | 3.55 | Identical | 874.4504 | 0.98 | 0.98 | Identical | On | On | Same | ¦12.030V 0.980A ¦ | ¦12.030V 0.980A ¦
180 | 879.3356 | 3548.9985 | 3.57 | 3.57 | Identical | 879.3356 | 0.98 | 0.98 | Identical | On | On | Same | ¦12.030V 0.980A ¦ | ¦12.030V 0.980A ¦
181 | 884.2208 | 3568.7152 | 3.59 | 3.58 | Within 10 mV (calibration line) | 884.2208 | 0.99 | 0.99 | Identical | On | On | Same | ¦12.030V 0.990A ¦ | ¦12.030V 0.990A ¦
182 | 889.106 | 3588.4319 | 3.61 | 3.6 | Within 10 mV (calibration line) | 889.106 | 1 | 1 | Identical | On | On | Same | ¦12.030V 1.000A ¦ | ¦12.030V 1.000A ¦
183 | 893.9912 | 3608.1485 | 3.63 | 3.62 | Within 10 mV (calibration line) | 893.9912 | 1 | 1 | Identical | On | On | Same | ¦12.030V 1.000A ¦ | ¦12.030V 1.000A ¦
184 | 898.8764 | 3627.8652 | 3.65 | 3.64 | Within 10 mV (calibration line) | 898.8764 | 1.01 | 1.01 | Identical | On | On | Same | ¦12.030V 1.010A ¦ | ¦12.030V 1.010A ¦
185 | 903.7616 | 3647.5818 | 3.67 | 3.66 | Within 10 mV (calibration line) | 903.7616 | 1.01 | 1.01 | Identical | On | On | Same | ¦12.030V 1.010A ¦ | ¦12.030V 1.010A ¦
186 | 908.6468 | 3667.2985 | 3.69 | 3.68 | Within 10 mV (calibration line) | 908.6468 | 1.02 | 1.02 | Identical | On | On | Same | ¦12.030V 1.020A ¦ | ¦12.030V 1.020A ¦
187 | 913.532 | 3687.0151 | 3.71 | 3.7 | Within 10 mV (calibration line) | 913.532 | 1.02 | 1.02 | Identical | On | On | Same | ¦12.030V 1.020A ¦ | ¦12.030V 1.020A ¦
188 | 918.4172 | 3706.7318 | 3.73 | 3.72 | Within 10 mV (calibration line) | 918.4172 | 1.03 | 1.03 | Identical | On | On | Same | ¦12.030V 1.030A ¦ | ¦12.030V 1.030A ¦
189 | 923.3024 | 3726.4485 | 3.75 | 3.74 | Within 10 mV (calibration line) | 923.3024 | 1.03 | 1.03 | Identical | On | On | Same | ¦12.030V 1.030A ¦ | ¦12.030V 1.030A ¦
190 | 928.1876 | 3746.1651 | 3.77 | 3.76 | Within 10 mV (calibration line) | 928.1876 | 1.04 | 1.04 | Identical | On | On | Same | ¦12.030V 1.040A ¦ | ¦12.030V 1.040A ¦
191 | 933.0728 | 3765.8818 | 3.79 | 3.78 | Within 10 mV (calibration line) | 933.0728 | 1.04 | 1.04 | Identical | On | On | Same | ¦12.030V 1.040A ¦ | ¦12.030V 1.040A ¦
192 | 937.958 | 3785.5984 | 3.81 | 3.8 | Within 10 mV (calibration line) | 937.958 | 1.05 | 1.05 | Identical | On | On | Same | ¦12.030V 1.050A ¦ | ¦12.030V 1.050A ¦
193 | 942.8432 | 3805.3151 | 3.83 | 3.82 | Within 10 mV (calibration line) | 942.8432 | 1.05 | 1.06 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 1.050A ¦ | ¦12.030V 1.060A ¦
194 | 947.7284 | 3825.0318 | 3.85 | 3.84 | Within 10 mV (calibration line) | 947.7284 | 1.06 | 1.06 | Identical | On | On | Same | ¦12.030V 1.060A ¦ | ¦12.030V 1.060A ¦
195 | 952.6136 | 3844.7484 | 3.87 | 3.86 | Within 10 mV (calibration line) | 952.6136 | 1.07 | 1.07 | Identical | On | On | Same | ¦12.030V 1.070A ¦ | ¦12.030V 1.070A ¦
196 | 957.4988 | 3864.4651 | 3.89 | 3.88 | Within 10 mV (calibration line) | 957.4988 | 1.07 | 1.07 | Identical | On | On | Same | ¦12.030V 1.070A ¦ | ¦12.030V 1.070A ¦
197 | 962.384 | 3884.1817 | 3.91 | 3.9 | Within 10 mV (calibration line) | 962.384 | 1.08 | 1.08 | Identical | On | On | Same | ¦12.030V 1.080A ¦ | ¦12.030V 1.080A ¦
198 | 967.2692 | 3903.8984 | 3.93 | 3.92 | Within 10 mV (calibration line) | 967.2692 | 1.08 | 1.08 | Identical | On | On | Same | ¦12.030V 1.080A ¦ | ¦12.030V 1.080A ¦
199 | 972.1544 | 3923.615 | 3.95 | 3.94 | Within 10 mV (calibration line) | 972.1544 | 1.09 | 1.09 | Identical | On | On | Same | ¦12.030V 1.090A ¦ | ¦12.030V 1.090A ¦
200 | 977.0396 | 3943.3317 | 3.97 | 3.96 | Within 10 mV (calibration line) | 977.0396 | 1.09 | 1.09 | Identical | On | On | Same | ¦12.030V 1.090A ¦ | ¦12.030V 1.090A ¦
201 | 981.9248 | 3963.0484 | 3.99 | 3.98 | Within 10 mV (calibration line) | 981.9248 | 1.1 | 1.1 | Identical | On | On | Same | ¦12.030V 1.100A ¦ | ¦12.030V 1.100A ¦
202 | 986.81 | 3982.765 | 4.01 | 4 | Within 10 mV (calibration line) | 986.81 | 1.1 | 1.11 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 1.100A ¦ | ¦12.030V 1.110A ¦
203 | 991.6952 | 4002.4817 | 4.03 | 4.02 | Within 10 mV (calibration line) | 991.6952 | 1.11 | 1.11 | Identical | On | On | Same | ¦12.030V 1.110A ¦ | ¦12.030V 1.110A ¦
204 | 996.5804 | 4022.1983 | 4.05 | 4.04 | Within 10 mV (calibration line) | 996.5804 | 1.12 | 1.12 | Identical | On | On | Same | ¦12.030V 1.120A ¦ | ¦12.030V 1.120A ¦
205 | 1001.4656 | 4041.915 | 4.07 | 4.06 | Within 10 mV (calibration line) | 1001.4656 | 1.12 | 1.12 | Identical | On | On | Same | ¦12.030V 1.120A ¦ | ¦12.030V 1.120A ¦
206 | 1006.3508 | 4061.6317 | 4.09 | 4.08 | Within 10 mV (calibration line) | 1006.3508 | 1.13 | 1.13 | Identical | On | On | Same | ¦12.030V 1.130A ¦ | ¦12.030V 1.130A ¦
207 | 1011.236 | 4081.3483 | 4.11 | 4.1 | Within 10 mV (calibration line) | 1011.236 | 1.13 | 1.13 | Identical | On | On | Same | ¦12.030V 1.130A ¦ | ¦12.030V 1.130A ¦
208 | 1016.1212 | 4101.065 | 4.13 | 4.11 | -20 mV (end of old +30 mV band) | 1016.1212 | 1.14 | 1.14 | Identical | On | On | Same | ¦12.030V 1.140A ¦ | ¦12.030V 1.140A ¦
209 | 1021.0064 | 4120.7816 | 4.15 | 4.13 | -20 mV (end of old +30 mV band) | 1021.0064 | 1.14 | 1.14 | Identical | On | On | Same | ¦12.030V 1.140A ¦ | ¦12.030V 1.140A ¦
210 | 1025.8915 | 4140.4983 | 4.17 | 4.15 | -20 mV (end of old +30 mV band) | 1025.8915 | 1.15 | 1.15 | Identical | On | On | Same | ¦12.030V 1.150A ¦ | ¦12.030V 1.150A ¦
211 | 1030.7767 | 4160.2149 | 4.17 | 4.17 | Identical | 1030.7767 | 1.15 | 1.15 | Identical | On | On | Same | ¦12.030V 1.150A ¦ | ¦12.030V 1.150A ¦
212 | 1035.6619 | 4179.9316 | 4.18 | 4.19 | Within 10 mV (calibration line) | 1035.6619 | 1.16 | 1.16 | Identical | On | On | Same | ¦12.030V 1.160A ¦ | ¦12.030V 1.160A ¦
213 | 1040.5471 | 4199.6483 | 4.2 | 4.21 | Within 10 mV (calibration line) | 1040.5471 | 1.16 | 1.17 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 1.160A ¦ | ¦12.030V 1.170A ¦
214 | 1045.4323 | 4219.3649 | 4.22 | 4.23 | Within 10 mV (calibration line) | 1045.4323 | 1.17 | 1.17 | Identical | On | On | Same | ¦12.030V 1.170A ¦ | ¦12.030V 1.170A ¦
215 | 1050.3175 | 4239.0816 | 4.24 | 4.25 | Within 10 mV (calibration line) | 1050.3175 | 1.18 | 1.18 | Identical | On | On | Same | ¦12.030V 1.180A ¦ | ¦12.030V 1.180A ¦
216 | 1055.2027 | 4258.7982 | 4.26 | 4.27 | Within 10 mV (calibration line) | 1055.2027 | 1.18 | 1.18 | Identical | On | On | Same | ¦12.030V 1.180A ¦ | ¦12.030V 1.180A ¦
217 | 1060.0879 | 4278.5149 | 4.28 | 4.29 | Within 10 mV (calibration line) | 1060.0879 | 1.19 | 1.19 | Identical | On | On | Same | ¦12.030V 1.190A ¦ | ¦12.030V 1.190A ¦
218 | 1064.9731 | 4298.2316 | 4.3 | 4.31 | Within 10 mV (calibration line) | 1064.9731 | 1.19 | 1.19 | Identical | On | On | Same | ¦12.030V 1.190A ¦ | ¦12.030V 1.190A ¦
219 | 1069.8583 | 4317.9482 | 4.32 | 4.33 | Within 10 mV (calibration line) | 1069.8583 | 1.2 | 1.2 | Identical | On | On | Same | ¦12.030V 1.200A ¦ | ¦12.030V 1.200A ¦
220 | 1074.7435 | 4337.6649 | 4.34 | 4.35 | Within 10 mV (calibration line) | 1074.7435 | 1.2 | 1.2 | Identical | On | On | Same | ¦12.030V 1.200A ¦ | ¦12.030V 1.200A ¦
221 | 1079.6287 | 4357.3815 | 4.36 | 4.37 | Within 10 mV (calibration line) | 1079.6287 | 1.21 | 1.21 | Identical | On | On | Same | ¦12.030V 1.210A ¦ | ¦12.030V 1.210A ¦
222 | 1084.5139 | 4377.0982 | 4.38 | 4.39 | Within 10 mV (calibration line) | 1084.5139 | 1.21 | 1.22 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 1.210A ¦ | ¦12.030V 1.220A ¦
223 | 1089.3991 | 4396.8149 | 4.4 | 4.41 | Within 10 mV (calibration line) | 1089.3991 | 1.22 | 1.22 | Identical | On | On | Same | ¦12.030V 1.220A ¦ | ¦12.030V 1.220A ¦
224 | 1094.2843 | 4416.5315 | 4.42 | 4.43 | Within 10 mV (calibration line) | 1094.2843 | 1.23 | 1.23 | Identical | On | On | Same | ¦12.030V 1.230A ¦ | ¦12.030V 1.230A ¦
225 | 1099.1695 | 4436.2482 | 4.44 | 4.45 | Within 10 mV (calibration line) | 1099.1695 | 1.23 | 1.23 | Identical | On | On | Same | ¦12.030V 1.230A ¦ | ¦12.030V 1.230A ¦
226 | 1104.0547 | 4455.9648 | 4.46 | 4.47 | Within 10 mV (calibration line) | 1104.0547 | 1.24 | 1.24 | Identical | On | On | Same | ¦12.030V 1.240A ¦ | ¦12.030V 1.240A ¦
227 | 1108.9399 | 4475.6815 | 4.48 | 4.49 | Within 10 mV (calibration line) | 1108.9399 | 1.24 | 1.24 | Identical | On | On | Same | ¦12.030V 1.240A ¦ | ¦12.030V 1.240A ¦
228 | 1113.8251 | 4495.3981 | 4.5 | 4.51 | Within 10 mV (calibration line) | 1113.8251 | 1.25 | 1.25 | Identical | On | On | Same | ¦12.030V 1.250A ¦ | ¦12.030V 1.250A ¦
229 | 1118.7103 | 4515.1148 | 4.52 | 4.53 | Within 10 mV (calibration line) | 1118.7103 | 1.25 | 1.25 | Identical | On | On | Same | ¦12.030V 1.250A ¦ | ¦12.030V 1.250A ¦
230 | 1123.5955 | 4534.8315 | 4.54 | 4.55 | Within 10 mV (calibration line) | 1123.5955 | 1.26 | 1.26 | Identical | On | On | Same | ¦12.030V 1.260A ¦ | ¦12.030V 1.260A ¦
231 | 1128.4807 | 4554.5481 | 4.56 | 4.57 | Within 10 mV (calibration line) | 1128.4807 | 1.26 | 1.26 | Identical | On | On | Same | ¦12.030V 1.260A ¦ | ¦12.030V 1.260A ¦
232 | 1133.3659 | 4574.2648 | 4.58 | 4.59 | Within 10 mV (calibration line) | 1133.3659 | 1.27 | 1.27 | Identical | On | On | Same | ¦12.030V 1.270A ¦ | ¦12.030V 1.270A ¦
233 | 1138.2511 | 4593.9814 | 4.6 | 4.61 | Within 10 mV (calibration line) | 1138.2511 | 1.28 | 1.28 | Identical | On | On | Same | ¦12.030V 1.280A ¦ | ¦12.030V 1.280A ¦
234 | 1143.1363 | 4613.6981 | 4.62 | 4.63 | Within 10 mV (calibration line) | 1143.1363 | 1.28 | 1.28 | Identical | On | On | Same | ¦12.030V 1.280A ¦ | ¦12.030V 1.280A ¦
235 | 1148.0215 | 4633.4148 | 4.64 | 4.64 | Identical | 1148.0215 | 1.29 | 1.29 | Identical | On | On | Same | ¦12.030V 1.290A ¦ | ¦12.030V 1.290A ¦
236 | 1152.9067 | 4653.1314 | 4.66 | 4.66 | Identical | 1152.9067 | 1.29 | 1.29 | Identical | On | On | Same | ¦12.030V 1.290A ¦ | ¦12.030V 1.290A ¦
237 | 1157.7919 | 4672.8481 | 4.68 | 4.68 | Identical | 1157.7919 | 1.3 | 1.3 | Identical | On | On | Same | ¦12.030V 1.300A ¦ | ¦12.030V 1.300A ¦
238 | 1162.6771 | 4692.5647 | 4.7 | 4.7 | Identical | 1162.6771 | 1.3 | 1.3 | Identical | On | On | Same | ¦12.030V 1.300A ¦ | ¦12.030V 1.300A ¦
239 | 1167.5623 | 4712.2814 | 4.72 | 4.72 | Identical | 1167.5623 | 1.31 | 1.31 | Identical | On | On | Same | ¦12.030V 1.310A ¦ | ¦12.030V 1.310A ¦
240 | 1172.4475 | 4731.998 | 4.74 | 4.74 | Identical | 1172.4475 | 1.31 | 1.31 | Identical | On | On | Same | ¦12.030V 1.310A ¦ | ¦12.030V 1.310A ¦
241 | 1177.3327 | 4751.7147 | 4.76 | 4.76 | Identical | 1177.3327 | 1.32 | 1.32 | Identical | On | On | Same | ¦12.030V 1.320A ¦ | ¦12.030V 1.320A ¦
242 | 1182.2179 | 4771.4314 | 4.78 | 4.78 | Identical | 1182.2179 | 1.32 | 1.32 | Identical | On | On | Same | ¦12.030V 1.320A ¦ | ¦12.030V 1.320A ¦
243 | 1187.1031 | 4791.148 | 4.8 | 4.8 | Identical | 1187.1031 | 1.33 | 1.33 | Identical | On | On | Same | ¦12.030V 1.330A ¦ | ¦12.030V 1.330A ¦
244 | 1191.9883 | 4810.8647 | 4.82 | 4.82 | Identical | 1191.9883 | 1.33 | 1.34 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 1.330A ¦ | ¦12.030V 1.340A ¦
245 | 1196.8735 | 4830.5813 | 4.84 | 4.84 | Identical | 1196.8735 | 1.34 | 1.34 | Identical | On | On | Same | ¦12.030V 1.340A ¦ | ¦12.030V 1.340A ¦
246 | 1201.7587 | 4850.298 | 4.86 | 4.86 | Identical | 1201.7587 | 1.35 | 1.35 | Identical | On | On | Same | ¦12.030V 1.350A ¦ | ¦12.030V 1.350A ¦
247 | 1206.6439 | 4870.0147 | 4.88 | 4.88 | Identical | 1206.6439 | 1.35 | 1.35 | Identical | On | On | Same | ¦12.030V 1.350A ¦ | ¦12.030V 1.350A ¦
248 | 1211.5291 | 4889.7313 | 4.89 | 4.9 | Within 10 mV (calibration line) | 1211.5291 | 1.36 | 1.36 | Identical | On | On | Same | ¦12.030V 1.360A ¦ | ¦12.030V 1.360A ¦
249 | 1216.4143 | 4909.448 | 4.91 | 4.92 | Within 10 mV (calibration line) | 1216.4143 | 1.36 | 1.36 | Identical | On | On | Same | ¦12.030V 1.360A ¦ | ¦12.030V 1.360A ¦
250 | 1221.2995 | 4929.1646 | 4.93 | 4.94 | Within 10 mV (calibration line) | 1221.2995 | 1.37 | 1.37 | Identical | On | On | Same | ¦12.030V 1.370A ¦ | ¦12.030V 1.370A ¦
251 | 1226.1847 | 4948.8813 | 4.95 | 4.96 | Within 10 mV (calibration line) | 1226.1847 | 1.37 | 1.37 | Identical | On | On | Same | ¦12.030V 1.370A ¦ | ¦12.030V 1.370A ¦
252 | 1231.0699 | 4968.5979 | 4.97 | 4.98 | Within 10 mV (calibration line) | 1231.0699 | 1.38 | 1.38 | Identical | On | On | Same | ¦12.030V 1.380A ¦ | ¦12.030V 1.380A ¦
253 | 1235.9551 | 4988.3146 | 4.99 | 5 | Within 10 mV (calibration line) | 1235.9551 | 1.38 | 1.39 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 1.380A ¦ | ¦12.030V 1.390A ¦
254 | 1240.8403 | 5008.0313 | 5.01 | 5.02 | Within 10 mV (calibration line) | 1240.8403 | 1.39 | 1.39 | Identical | On | On | Same | ¦12.030V 1.390A ¦ | ¦12.030V 1.390A ¦
255 | 1245.7255 | 5027.7479 | 5.03 | 5.04 | Within 10 mV (calibration line) | 1245.7255 | 1.4 | 1.4 | Identical | On | On | Same | ¦12.030V 1.400A ¦ | ¦12.030V 1.400A ¦
256 | 1250.6106 | 5047.4646 | 5.05 | 5.06 | Within 10 mV (calibration line) | 1250.6106 | 1.4 | 1.4 | Identical | On | On | Same | ¦12.030V 1.400A ¦ | ¦12.030V 1.400A ¦
257 | 1255.4958 | 5067.1812 | 5.07 | 5.08 | Within 10 mV (calibration line) | 1255.4958 | 1.41 | 1.41 | Identical | On | On | Same | ¦12.030V 1.410A ¦ | ¦12.030V 1.410A ¦
258 | 1260.381 | 5086.8979 | 5.09 | 5.1 | Within 10 mV (calibration line) | 1260.381 | 1.41 | 1.41 | Identical | On | On | Same | ¦12.030V 1.410A ¦ | ¦12.030V 1.410A ¦
259 | 1265.2662 | 5106.6146 | 5.11 | 5.12 | Within 10 mV (calibration line) | 1265.2662 | 1.42 | 1.42 | Identical | On | On | Same | ¦12.030V 1.420A ¦ | ¦12.030V 1.420A ¦
260 | 1270.1514 | 5126.3312 | 5.13 | 5.14 | Within 10 mV (calibration line) | 1270.1514 | 1.42 | 1.42 | Identical | On | On | Same | ¦12.030V 1.420A ¦ | ¦12.030V 1.420A ¦
261 | 1275.0366 | 5146.0479 | 5.15 | 5.16 | Within 10 mV (calibration line) | 1275.0366 | 1.43 | 1.43 | Identical | On | On | Same | ¦12.030V 1.430A ¦ | ¦12.030V 1.430A ¦
262 | 1279.9218 | 5165.7645 | 5.17 | 5.17 | Identical | 1279.9218 | 1.43 | 1.44 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 1.430A ¦ | ¦12.030V 1.440A ¦
263 | 1284.807 | 5185.4812 | 5.19 | 5.19 | Identical | 1284.807 | 1.44 | 1.44 | Identical | On | On | Same | ¦12.030V 1.440A ¦ | ¦12.030V 1.440A ¦
264 | 1289.6922 | 5205.1979 | 5.21 | 5.21 | Identical | 1289.6922 | 1.45 | 1.45 | Identical | On | On | Same | ¦12.030V 1.450A ¦ | ¦12.030V 1.450A ¦
265 | 1294.5774 | 5224.9145 | 5.23 | 5.23 | Identical | 1294.5774 | 1.45 | 1.45 | Identical | On | On | Same | ¦12.030V 1.450A ¦ | ¦12.030V 1.450A ¦
266 | 1299.4626 | 5244.6312 | 5.25 | 5.25 | Identical | 1299.4626 | 1.46 | 1.46 | Identical | On | On | Same | ¦12.030V 1.460A ¦ | ¦12.030V 1.460A ¦
267 | 1304.3478 | 5264.3478 | 5.27 | 5.27 | Identical | 1304.3478 | 1.46 | 1.46 | Identical | On | On | Same | ¦12.030V 1.460A ¦ | ¦12.030V 1.460A ¦
268 | 1309.233 | 5284.0645 | 5.29 | 5.29 | Identical | 1309.233 | 1.47 | 1.47 | Identical | On | On | Same | ¦12.030V 1.470A ¦ | ¦12.030V 1.470A ¦
269 | 1314.1182 | 5303.7811 | 5.31 | 5.31 | Identical | 1314.1182 | 1.47 | 1.47 | Identical | On | On | Same | ¦12.030V 1.470A ¦ | ¦12.030V 1.470A ¦
270 | 1319.0034 | 5323.4978 | 5.33 | 5.33 | Identical | 1319.0034 | 1.48 | 1.48 | Identical | On | On | Same | ¦12.030V 1.480A ¦ | ¦12.030V 1.480A ¦
271 | 1323.8886 | 5343.2145 | 5.35 | 5.35 | Identical | 1323.8886 | 1.48 | 1.48 | Identical | On | On | Same | ¦12.030V 1.480A ¦ | ¦12.030V 1.480A ¦
272 | 1328.7738 | 5362.9311 | 5.37 | 5.37 | Identical | 1328.7738 | 1.49 | 1.49 | Identical | On | On | Same | ¦12.030V 1.490A ¦ | ¦12.030V 1.490A ¦
273 | 1333.659 | 5382.6478 | 5.39 | 5.39 | Identical | 1333.659 | 1.49 | 1.5 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 1.490A ¦ | ¦12.030V 1.500A ¦
274 | 1338.5442 | 5402.3644 | 5.41 | 5.41 | Identical | 1338.5442 | 1.5 | 1.5 | Identical | On | On | Same | ¦12.030V 1.500A ¦ | ¦12.030V 1.500A ¦
275 | 1343.4294 | 5422.0811 | 5.43 | 5.43 | Identical | 1343.4294 | 1.51 | 1.51 | Identical | On | On | Same | ¦12.030V 1.510A ¦ | ¦12.030V 1.510A ¦
276 | 1348.3146 | 5441.7978 | 5.45 | 5.45 | Identical | 1348.3146 | 1.51 | 1.51 | Identical | On | On | Same | ¦12.030V 1.510A ¦ | ¦12.030V 1.510A ¦
277 | 1353.1998 | 5461.5144 | 5.47 | 5.47 | Identical | 1353.1998 | 1.52 | 1.52 | Identical | On | On | Same | ¦12.030V 1.520A ¦ | ¦12.030V 1.520A ¦
278 | 1358.085 | 5481.2311 | 5.49 | 5.49 | Identical | 1358.085 | 1.52 | 1.52 | Identical | On | On | Same | ¦12.030V 1.520A ¦ | ¦12.030V 1.520A ¦
279 | 1362.9702 | 5500.9477 | 5.51 | 5.51 | Identical | 1362.9702 | 1.53 | 1.53 | Identical | On | On | Same | ¦12.030V 1.530A ¦ | ¦12.030V 1.530A ¦
280 | 1367.8554 | 5520.6644 | 5.53 | 5.53 | Identical | 1367.8554 | 1.53 | 1.53 | Identical | On | On | Same | ¦12.030V 1.530A ¦ | ¦12.030V 1.530A ¦
281 | 1372.7406 | 5540.381 | 5.55 | 5.55 | Identical | 1372.7406 | 1.54 | 1.54 | Identical | On | On | Same | ¦12.030V 1.540A ¦ | ¦12.030V 1.540A ¦
282 | 1377.6258 | 5560.0977 | 5.57 | 5.57 | Identical | 1377.6258 | 1.54 | 1.54 | Identical | On | On | Same | ¦12.030V 1.540A ¦ | ¦12.030V 1.540A ¦
283 | 1382.511 | 5579.8144 | 5.58 | 5.59 | Within 10 mV (calibration line) | 1382.511 | 1.55 | 1.55 | Identical | On | On | Same | ¦12.030V 1.550A ¦ | ¦12.030V 1.550A ¦
284 | 1387.3962 | 5599.531 | 5.6 | 5.61 | Within 10 mV (calibration line) | 1387.3962 | 1.55 | 1.55 | Identical | On | On | Same | ¦12.030V 1.550A ¦ | ¦12.030V 1.550A ¦
285 | 1392.2814 | 5619.2477 | 5.62 | 5.63 | Within 10 mV (calibration line) | 1392.2814 | 1.56 | 1.56 | Identical | On | On | Same | ¦12.030V 1.560A ¦ | ¦12.030V 1.560A ¦
286 | 1397.1666 | 5638.9643 | 5.64 | 5.65 | Within 10 mV (calibration line) | 1397.1666 | 1.57 | 1.57 | Identical | On | On | Same | ¦12.030V 1.570A ¦ | ¦12.030V 1.570A ¦
287 | 1402.0518 | 5658.681 | 5.66 | 5.67 | Within 10 mV (calibration line) | 1402.0518 | 1.57 | 1.57 | Identical | On | On | Same | ¦12.030V 1.570A ¦ | ¦12.030V 1.570A ¦
288 | 1406.937 | 5678.3977 | 5.68 | 5.69 | Within 10 mV (calibration line) | 1406.937 | 1.58 | 1.58 | Identical | On | On | Same | ¦12.030V 1.580A ¦ | ¦12.030V 1.580A ¦
289 | 1411.8222 | 5698.1143 | 5.7 | 5.71 | Within 10 mV (calibration line) | 1411.8222 | 1.58 | 1.58 | Identical | On | On | Same | ¦12.030V 1.580A ¦ | ¦12.030V 1.580A ¦
290 | 1416.7074 | 5717.831 | 5.72 | 5.72 | Identical | 1416.7074 | 1.59 | 1.59 | Identical | On | On | Same | ¦12.030V 1.590A ¦ | ¦12.030V 1.590A ¦
291 | 1421.5926 | 5737.5476 | 5.74 | 5.74 | Identical | 1421.5926 | 1.59 | 1.59 | Identical | On | On | Same | ¦12.030V 1.590A ¦ | ¦12.030V 1.590A ¦
292 | 1426.4778 | 5757.2643 | 5.76 | 5.76 | Identical | 1426.4778 | 1.6 | 1.6 | Identical | On | On | Same | ¦12.030V 1.600A ¦ | ¦12.030V 1.600A ¦
293 | 1431.363 | 5776.9809 | 5.78 | 5.78 | Identical | 1431.363 | 1.6 | 1.6 | Identical | On | On | Same | ¦12.030V 1.600A ¦ | ¦12.030V 1.600A ¦
294 | 1436.2482 | 5796.6976 | 5.8 | 5.8 | Identical | 1436.2482 | 1.61 | 1.61 | Identical | On | On | Same | ¦12.030V 1.610A ¦ | ¦12.030V 1.610A ¦
295 | 1441.1334 | 5816.4143 | 5.82 | 5.82 | Identical | 1441.1334 | 1.62 | 1.62 | Identical | On | On | Same | ¦12.030V 1.620A ¦ | ¦12.030V 1.620A ¦
296 | 1446.0186 | 5836.1309 | 5.84 | 5.84 | Identical | 1446.0186 | 1.62 | 1.62 | Identical | On | On | Same | ¦12.030V 1.620A ¦ | ¦12.030V 1.620A ¦
297 | 1450.9038 | 5855.8476 | 5.86 | 5.86 | Identical | 1450.9038 | 1.63 | 1.63 | Identical | On | On | Same | ¦12.030V 1.630A ¦ | ¦12.030V 1.630A ¦
298 | 1455.789 | 5875.5642 | 5.88 | 5.88 | Identical | 1455.789 | 1.63 | 1.63 | Identical | On | On | Same | ¦12.030V 1.630A ¦ | ¦12.030V 1.630A ¦
299 | 1460.6742 | 5895.2809 | 5.9 | 5.9 | Identical | 1460.6742 | 1.64 | 1.64 | Identical | On | On | Same | ¦12.030V 1.640A ¦ | ¦12.030V 1.640A ¦
300 | 1465.5594 | 5914.9976 | 5.92 | 5.92 | Identical | 1465.5594 | 1.64 | 1.64 | Identical | On | On | Same | ¦12.030V 1.640A ¦ | ¦12.030V 1.640A ¦
301 | 1470.4446 | 5934.7142 | 5.94 | 5.94 | Identical | 1470.4446 | 1.65 | 1.65 | Identical | On | On | Same | ¦12.030V 1.650A ¦ | ¦12.030V 1.650A ¦
302 | 1475.3298 | 5954.4309 | 5.96 | 5.96 | Identical | 1475.3298 | 1.65 | 1.65 | Identical | On | On | Same | ¦12.030V 1.650A ¦ | ¦12.030V 1.650A ¦
303 | 1480.2149 | 5974.1475 | 5.98 | 5.98 | Identical | 1480.2149 | 1.66 | 1.66 | Identical | On | On | Same | ¦12.030V 1.660A ¦ | ¦12.030V 1.660A ¦
304 | 1485.1001 | 5993.8642 | 5.99 | 6 | Within 10 mV (calibration line) | 1485.1001 | 1.67 | 1.67 | Identical | On | On | Same | ¦12.030V 1.670A ¦ | ¦12.030V 1.670A ¦
305 | 1489.9853 | 6013.5809 | 6.01 | 6.02 | Within 10 mV (calibration line) | 1489.9853 | 1.67 | 1.67 | Identical | On | On | Same | ¦12.030V 1.670A ¦ | ¦12.030V 1.670A ¦
306 | 1494.8705 | 6033.2975 | 6.03 | 6.04 | Within 10 mV (calibration line) | 1494.8705 | 1.68 | 1.68 | Identical | On | On | Same | ¦12.030V 1.680A ¦ | ¦12.030V 1.680A ¦
307 | 1499.7557 | 6053.0142 | 6.05 | 6.06 | Within 10 mV (calibration line) | 1499.7557 | 1.68 | 1.68 | Identical | On | On | Same | ¦12.030V 1.680A ¦ | ¦12.030V 1.680A ¦
308 | 1504.6409 | 6072.7308 | 6.07 | 6.08 | Within 10 mV (calibration line) | 1504.6409 | 1.69 | 1.69 | Identical | On | On | Same | ¦12.030V 1.690A ¦ | ¦12.030V 1.690A ¦
309 | 1509.5261 | 6092.4475 | 6.09 | 6.1 | Within 10 mV (calibration line) | 1509.5261 | 1.69 | 1.69 | Identical | On | On | Same | ¦12.030V 1.690A ¦ | ¦12.030V 1.690A ¦
310 | 1514.4113 | 6112.1641 | 6.11 | 6.12 | Within 10 mV (calibration line) | 1514.4113 | 1.7 | 1.7 | Identical | On | On | Same | ¦12.030V 1.700A ¦ | ¦12.030V 1.700A ¦
311 | 1519.2965 | 6131.8808 | 6.13 | 6.14 | Within 10 mV (calibration line) | 1519.2965 | 1.7 | 1.7 | Identical | On | On | Same | ¦12.030V 1.700A ¦ | ¦12.030V 1.700A ¦
312 | 1524.1817 | 6151.5975 | 6.15 | 6.16 | Within 10 mV (calibration line) | 1524.1817 | 1.71 | 1.71 | Identical | On | On | Same | ¦12.030V 1.710A ¦ | ¦12.030V 1.710A ¦
313 | 1529.0669 | 6171.3141 | 6.17 | 6.18 | Within 10 mV (calibration line) | 1529.0669 | 1.71 | 1.71 | Identical | On | On | Same | ¦12.030V 1.710A ¦ | ¦12.030V 1.710A ¦
314 | 1533.9521 | 6191.0308 | 6.19 | 6.2 | Within 10 mV (calibration line) | 1533.9521 | 1.72 | 1.72 | Identical | On | On | Same | ¦12.030V 1.720A ¦ | ¦12.030V 1.720A ¦
315 | 1538.8373 | 6210.7474 | 6.21 | 6.22 | Within 10 mV (calibration line) | 1538.8373 | 1.72 | 1.73 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 1.720A ¦ | ¦12.030V 1.730A ¦
316 | 1543.7225 | 6230.4641 | 6.23 | 6.24 | Within 10 mV (calibration line) | 1543.7225 | 1.73 | 1.73 | Identical | On | On | Same | ¦12.030V 1.730A ¦ | ¦12.030V 1.730A ¦
317 | 1548.6077 | 6250.1808 | 6.25 | 6.25 | Identical | 1548.6077 | 1.74 | 1.74 | Identical | On | On | Same | ¦12.030V 1.740A ¦ | ¦12.030V 1.740A ¦
318 | 1553.4929 | 6269.8974 | 6.26 | 6.27 | Within 10 mV (calibration line) | 1553.4929 | 1.74 | 1.74 | Identical | On | On | Same | ¦12.030V 1.740A ¦ | ¦12.030V 1.740A ¦
319 | 1558.3781 | 6289.6141 | 6.28 | 6.29 | Within 10 mV (calibration line) | 1558.3781 | 1.75 | 1.75 | Identical | On | On | Same | ¦12.030V 1.750A ¦ | ¦12.030V 1.750A ¦
320 | 1563.2633 | 6309.3307 | 6.3 | 6.31 | Within 10 mV (calibration line) | 1563.2633 | 1.75 | 1.75 | Identical | On | On | Same | ¦12.030V 1.750A ¦ | ¦12.030V 1.750A ¦
321 | 1568.1485 | 6329.0474 | 6.32 | 6.33 | Within 10 mV (calibration line) | 1568.1485 | 1.76 | 1.76 | Identical | On | On | Same | ¦12.030V 1.760A ¦ | ¦12.030V 1.760A ¦
322 | 1573.0337 | 6348.764 | 6.34 | 6.35 | Within 10 mV (calibration line) | 1573.0337 | 1.76 | 1.76 | Identical | On | On | Same | ¦12.030V 1.760A ¦ | ¦12.030V 1.760A ¦
323 | 1577.9189 | 6368.4807 | 6.36 | 6.37 | Within 10 mV (calibration line) | 1577.9189 | 1.77 | 1.77 | Identical | On | On | Same | ¦12.030V 1.770A ¦ | ¦12.030V 1.770A ¦
324 | 1582.8041 | 6388.1974 | 6.38 | 6.39 | Within 10 mV (calibration line) | 1582.8041 | 1.77 | 1.77 | Identical | On | On | Same | ¦12.030V 1.770A ¦ | ¦12.030V 1.770A ¦
325 | 1587.6893 | 6407.914 | 6.4 | 6.41 | Within 10 mV (calibration line) | 1587.6893 | 1.78 | 1.78 | Identical | On | On | Same | ¦12.030V 1.780A ¦ | ¦12.030V 1.780A ¦
326 | 1592.5745 | 6427.6307 | 6.42 | 6.43 | Within 10 mV (calibration line) | 1592.5745 | 1.78 | 1.79 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 1.780A ¦ | ¦12.030V 1.790A ¦
327 | 1597.4597 | 6447.3473 | 6.44 | 6.45 | Within 10 mV (calibration line) | 1597.4597 | 1.79 | 1.79 | Identical | On | On | Same | ¦12.030V 1.790A ¦ | ¦12.030V 1.790A ¦
328 | 1602.3449 | 6467.064 | 6.46 | 6.47 | Within 10 mV (calibration line) | 1602.3449 | 1.8 | 1.8 | Identical | On | On | Same | ¦12.030V 1.800A ¦ | ¦12.030V 1.800A ¦
329 | 1607.2301 | 6486.7807 | 6.48 | 6.49 | Within 10 mV (calibration line) | 1607.2301 | 1.8 | 1.8 | Identical | On | On | Same | ¦12.030V 1.800A ¦ | ¦12.030V 1.800A ¦
330 | 1612.1153 | 6506.4973 | 6.5 | 6.51 | Within 10 mV (calibration line) | 1612.1153 | 1.81 | 1.81 | Identical | On | On | Same | ¦12.030V 1.810A ¦ | ¦12.030V 1.810A ¦
331 | 1617.0005 | 6526.214 | 6.52 | 6.53 | Within 10 mV (calibration line) | 1617.0005 | 1.81 | 1.81 | Identical | On | On | Same | ¦12.030V 1.810A ¦ | ¦12.030V 1.810A ¦
332 | 1621.8857 | 6545.9306 | 6.54 | 6.55 | Within 10 mV (calibration line) | 1621.8857 | 1.82 | 1.82 | Identical | On | On | Same | ¦12.030V 1.820A ¦ | ¦12.030V 1.820A ¦
333 | 1626.7709 | 6565.6473 | 6.56 | 6.57 | Within 10 mV (calibration line) | 1626.7709 | 1.82 | 1.82 | Identical | On | On | Same | ¦12.030V 1.820A ¦ | ¦12.030V 1.820A ¦
334 | 1631.6561 | 6585.3639 | 6.58 | 6.59 | Within 10 mV (calibration line) | 1631.6561 | 1.83 | 1.83 | Identical | On | On | Same | ¦12.030V 1.830A ¦ | ¦12.030V 1.830A ¦
335 | 1636.5413 | 6605.0806 | 6.6 | 6.61 | Within 10 mV (calibration line) | 1636.5413 | 1.83 | 1.84 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 1.830A ¦ | ¦12.030V 1.840A ¦
336 | 1641.4265 | 6624.7973 | 6.62 | 6.63 | Within 10 mV (calibration line) | 1641.4265 | 1.84 | 1.84 | Identical | On | On | Same | ¦12.030V 1.840A ¦ | ¦12.030V 1.840A ¦
337 | 1646.3117 | 6644.5139 | 6.64 | 6.65 | Within 10 mV (calibration line) | 1646.3117 | 1.85 | 1.85 | Identical | On | On | Same | ¦12.030V 1.850A ¦ | ¦12.030V 1.850A ¦
338 | 1651.1969 | 6664.2306 | 6.66 | 6.67 | Within 10 mV (calibration line) | 1651.1969 | 1.85 | 1.85 | Identical | On | On | Same | ¦12.030V 1.850A ¦ | ¦12.030V 1.850A ¦
339 | 1656.0821 | 6683.9472 | 6.68 | 6.69 | Within 10 mV (calibration line) | 1656.0821 | 1.86 | 1.86 | Identical | On | On | Same | ¦12.030V 1.860A ¦ | ¦12.030V 1.860A ¦
340 | 1660.9673 | 6703.6639 | 6.7 | 6.71 | Within 10 mV (calibration line) | 1660.9673 | 1.86 | 1.86 | Identical | On | On | Same | ¦12.030V 1.860A ¦ | ¦12.030V 1.860A ¦
341 | 1665.8525 | 6723.3806 | 6.72 | 6.73 | Within 10 mV (calibration line) | 1665.8525 | 1.87 | 1.87 | Identical | On | On | Same | ¦12.030V 1.870A ¦ | ¦12.030V 1.870A ¦
342 | 1670.7377 | 6743.0972 | 6.74 | 6.75 | Within 10 mV (calibration line) | 1670.7377 | 1.87 | 1.87 | Identical | On | On | Same | ¦12.030V 1.870A ¦ | ¦12.030V 1.870A ¦
343 | 1675.6229 | 6762.8139 | 6.76 | 6.77 | Within 10 mV (calibration line) | 1675.6229 | 1.88 | 1.88 | Identical | On | On | Same | ¦12.030V 1.880A ¦ | ¦12.030V 1.880A ¦
344 | 1680.5081 | 6782.5305 | 6.78 | 6.78 | Identical | 1680.5081 | 1.88 | 1.88 | Identical | On | On | Same | ¦12.030V 1.880A ¦ | ¦12.030V 1.880A ¦
345 | 1685.3933 | 6802.2472 | 6.8 | 6.8 | Identical | 1685.3933 | 1.89 | 1.89 | Identical | On | On | Same | ¦12.030V 1.890A ¦ | ¦12.030V 1.890A ¦
346 | 1690.2785 | 6821.9638 | 6.82 | 6.82 | Identical | 1690.2785 | 1.89 | 1.89 | Identical | On | On | Same | ¦12.030V 1.890A ¦ | ¦12.030V 1.890A ¦
347 | 1695.1637 | 6841.6805 | 6.84 | 6.84 | Identical | 1695.1637 | 1.9 | 1.9 | Identical | On | On | Same | ¦12.030V 1.900A ¦ | ¦12.030V 1.900A ¦
348 | 1700.0489 | 6861.3972 | 6.86 | 6.86 | Identical | 1700.0489 | 1.91 | 1.91 | Identical | On | On | Same | ¦12.030V 1.910A ¦ | ¦12.030V 1.910A ¦
349 | 1704.934 | 6881.1138 | 6.88 | 6.88 | Identical | 1704.934 | 1.91 | 1.91 | Identical | On | On | Same | ¦12.030V 1.910A ¦ | ¦12.030V 1.910A ¦
350 | 1709.8192 | 6900.8305 | 6.9 | 6.9 | Identical | 1709.8192 | 1.92 | 1.92 | Identical | On | On | Same | ¦12.030V 1.920A ¦ | ¦12.030V 1.920A ¦
351 | 1714.7044 | 6920.5471 | 6.92 | 6.92 | Identical | 1714.7044 | 1.92 | 1.92 | Identical | On | On | Same | ¦12.030V 1.920A ¦ | ¦12.030V 1.920A ¦
352 | 1719.5896 | 6940.2638 | 6.94 | 6.94 | Identical | 1719.5896 | 1.93 | 1.93 | Identical | On | On | Same | ¦12.030V 1.930A ¦ | ¦12.030V 1.930A ¦
353 | 1724.4748 | 6959.9805 | 6.95 | 6.96 | Within 10 mV (calibration line) | 1724.4748 | 1.93 | 1.93 | Identical | On | On | Same | ¦12.030V 1.930A ¦ | ¦12.030V 1.930A ¦
354 | 1729.36 | 6979.6971 | 6.97 | 6.98 | Within 10 mV (calibration line) | 1729.36 | 1.94 | 1.94 | Identical | On | On | Same | ¦12.030V 1.940A ¦ | ¦12.030V 1.940A ¦
355 | 1734.2452 | 6999.4138 | 6.99 | 7 | Within 10 mV (calibration line) | 1734.2452 | 1.94 | 1.94 | Identical | On | On | Same | ¦12.030V 1.940A ¦ | ¦12.030V 1.940A ¦
356 | 1739.1304 | 7019.1304 | 7.01 | 7.02 | Within 10 mV (calibration line) | 1739.1304 | 1.95 | 1.95 | Identical | On | On | Same | ¦12.030V 1.950A ¦ | ¦12.030V 1.950A ¦
357 | 1744.0156 | 7038.8471 | 7.03 | 7.04 | Within 10 mV (calibration line) | 1744.0156 | 1.96 | 1.96 | Identical | On | On | Same | ¦12.030V 1.960A ¦ | ¦12.030V 1.960A ¦
358 | 1748.9008 | 7058.5638 | 7.05 | 7.06 | Within 10 mV (calibration line) | 1748.9008 | 1.96 | 1.96 | Identical | On | On | Same | ¦12.030V 1.960A ¦ | ¦12.030V 1.960A ¦
359 | 1753.786 | 7078.2804 | 7.07 | 7.08 | Within 10 mV (calibration line) | 1753.786 | 1.97 | 1.97 | Identical | On | On | Same | ¦12.030V 1.970A ¦ | ¦12.030V 1.970A ¦
360 | 1758.6712 | 7097.9971 | 7.09 | 7.1 | Within 10 mV (calibration line) | 1758.6712 | 1.97 | 1.97 | Identical | On | On | Same | ¦12.030V 1.970A ¦ | ¦12.030V 1.970A ¦
361 | 1763.5564 | 7117.7137 | 7.11 | 7.12 | Within 10 mV (calibration line) | 1763.5564 | 1.98 | 1.98 | Identical | On | On | Same | ¦12.030V 1.980A ¦ | ¦12.030V 1.980A ¦
362 | 1768.4416 | 7137.4304 | 7.13 | 7.14 | Within 10 mV (calibration line) | 1768.4416 | 1.98 | 1.98 | Identical | On | On | Same | ¦12.030V 1.980A ¦ | ¦12.030V 1.980A ¦
363 | 1773.3268 | 7157.147 | 7.15 | 7.16 | Within 10 mV (calibration line) | 1773.3268 | 1.99 | 1.99 | Identical | On | On | Same | ¦12.030V 1.990A ¦ | ¦12.030V 1.990A ¦
364 | 1778.212 | 7176.8637 | 7.17 | 7.18 | Within 10 mV (calibration line) | 1778.212 | 1.99 | 1.99 | Identical | On | On | Same | ¦12.030V 1.990A ¦ | ¦12.030V 1.990A ¦
365 | 1783.0972 | 7196.5804 | 7.19 | 7.2 | Within 10 mV (calibration line) | 1783.0972 | 2 | 2 | Identical | On | On | Same | ¦12.030V 2.000A ¦ | ¦12.030V 2.000A ¦
366 | 1787.9824 | 7216.297 | 7.21 | 7.22 | Within 10 mV (calibration line) | 1787.9824 | 2 | 2.01 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 2.000A ¦ | ¦12.030V 2.010A ¦
367 | 1792.8676 | 7236.0137 | 7.23 | 7.24 | Within 10 mV (calibration line) | 1792.8676 | 2.01 | 2.01 | Identical | On | On | Same | ¦12.030V 2.010A ¦ | ¦12.030V 2.010A ¦
368 | 1797.7528 | 7255.7303 | 7.25 | 7.26 | Within 10 mV (calibration line) | 1797.7528 | 2.02 | 2.02 | Identical | On | On | Same | ¦12.030V 2.020A ¦ | ¦12.030V 2.020A ¦
369 | 1802.638 | 7275.447 | 7.27 | 7.28 | Within 10 mV (calibration line) | 1802.638 | 2.02 | 2.02 | Identical | On | On | Same | ¦12.030V 2.020A ¦ | ¦12.030V 2.020A ¦
370 | 1807.5232 | 7295.1637 | 7.29 | 7.3 | Within 10 mV (calibration line) | 1807.5232 | 2.03 | 2.03 | Identical | On | On | Same | ¦12.030V 2.030A ¦ | ¦12.030V 2.030A ¦
371 | 1812.4084 | 7314.8803 | 7.31 | 7.32 | Within 10 mV (calibration line) | 1812.4084 | 2.03 | 2.03 | Identical | On | On | Same | ¦12.030V 2.030A ¦ | ¦12.030V 2.030A ¦
372 | 1817.2936 | 7334.597 | 7.33 | 7.33 | Identical | 1817.2936 | 2.04 | 2.04 | Identical | On | On | Same | ¦12.030V 2.040A ¦ | ¦12.030V 2.040A ¦
373 | 1822.1788 | 7354.3136 | 7.35 | 7.35 | Identical | 1822.1788 | 2.04 | 2.04 | Identical | On | On | Same | ¦12.030V 2.040A ¦ | ¦12.030V 2.040A ¦
374 | 1827.064 | 7374.0303 | 7.37 | 7.37 | Identical | 1827.064 | 2.05 | 2.05 | Identical | On | On | Same | ¦12.030V 2.050A ¦ | ¦12.030V 2.050A ¦
375 | 1831.9492 | 7393.7469 | 7.39 | 7.39 | Identical | 1831.9492 | 2.05 | 2.05 | Identical | On | On | Same | ¦12.030V 2.050A ¦ | ¦12.030V 2.050A ¦
376 | 1836.8344 | 7413.4636 | 7.41 | 7.41 | Identical | 1836.8344 | 2.06 | 2.06 | Identical | On | On | Same | ¦12.030V 2.060A ¦ | ¦12.030V 2.060A ¦
377 | 1841.7196 | 7433.1803 | 7.43 | 7.43 | Identical | 1841.7196 | 2.06 | 2.07 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 2.060A ¦ | ¦12.030V 2.070A ¦
378 | 1846.6048 | 7452.8969 | 7.45 | 7.45 | Identical | 1846.6048 | 2.07 | 2.07 | Identical | On | On | Same | ¦12.030V 2.070A ¦ | ¦12.030V 2.070A ¦
379 | 1851.49 | 7472.6136 | 7.47 | 7.47 | Identical | 1851.49 | 2.08 | 2.08 | Identical | On | On | Same | ¦12.030V 2.080A ¦ | ¦12.030V 2.080A ¦
380 | 1856.3752 | 7492.3302 | 7.49 | 7.49 | Identical | 1856.3752 | 2.08 | 2.08 | Identical | On | On | Same | ¦12.030V 2.080A ¦ | ¦12.030V 2.080A ¦
381 | 1861.2604 | 7512.0469 | 7.51 | 7.51 | Identical | 1861.2604 | 2.09 | 2.09 | Identical | On | On | Same | ¦12.030V 2.090A ¦ | ¦12.030V 2.090A ¦
382 | 1866.1456 | 7531.7636 | 7.53 | 7.53 | Identical | 1866.1456 | 2.09 | 2.09 | Identical | On | On | Same | ¦12.030V 2.090A ¦ | ¦12.030V 2.090A ¦
383 | 1871.0308 | 7551.4802 | 7.55 | 7.55 | Identical | 1871.0308 | 2.1 | 2.1 | Identical | On | On | Same | ¦12.030V 2.100A ¦ | ¦12.030V 2.100A ¦
384 | 1875.916 | 7571.1969 | 7.57 | 7.57 | Identical | 1875.916 | 2.1 | 2.1 | Identical | On | On | Same | ¦12.030V 2.100A ¦ | ¦12.030V 2.100A ¦
385 | 1880.8012 | 7590.9135 | 7.59 | 7.59 | Identical | 1880.8012 | 2.11 | 2.11 | Identical | On | On | Same | ¦12.030V 2.110A ¦ | ¦12.030V 2.110A ¦
386 | 1885.6864 | 7610.6302 | 7.61 | 7.61 | Identical | 1885.6864 | 2.11 | 2.11 | Identical | On | On | Same | ¦12.030V 2.110A ¦ | ¦12.030V 2.110A ¦
387 | 1890.5716 | 7630.3468 | 7.63 | 7.63 | Identical | 1890.5716 | 2.12 | 2.12 | Identical | On | On | Same | ¦12.030V 2.120A ¦ | ¦12.030V 2.120A ¦
388 | 1895.4568 | 7650.0635 | 7.65 | 7.65 | Identical | 1895.4568 | 2.13 | 2.13 | Identical | On | On | Same | ¦12.030V 2.130A ¦ | ¦12.030V 2.130A ¦
389 | 1900.342 | 7669.7802 | 7.66 | 7.67 | Within 10 mV (calibration line) | 1900.342 | 2.13 | 2.13 | Identical | On | On | Same | ¦12.030V 2.130A ¦ | ¦12.030V 2.130A ¦
390 | 1905.2272 | 7689.4968 | 7.68 | 7.69 | Within 10 mV (calibration line) | 1905.2272 | 2.14 | 2.14 | Identical | On | On | Same | ¦12.030V 2.140A ¦ | ¦12.030V 2.140A ¦
391 | 1910.1124 | 7709.2135 | 7.7 | 7.71 | Within 10 mV (calibration line) | 1910.1124 | 2.14 | 2.14 | Identical | On | On | Same | ¦12.030V 2.140A ¦ | ¦12.030V 2.140A ¦
392 | 1914.9976 | 7728.9301 | 7.72 | 7.73 | Within 10 mV (calibration line) | 1914.9976 | 2.15 | 2.15 | Identical | On | On | Same | ¦12.030V 2.150A ¦ | ¦12.030V 2.150A ¦
393 | 1919.8828 | 7748.6468 | 7.74 | 7.75 | Within 10 mV (calibration line) | 1919.8828 | 2.15 | 2.15 | Identical | On | On | Same | ¦12.030V 2.150A ¦ | ¦12.030V 2.150A ¦
394 | 1924.768 | 7768.3635 | 7.76 | 7.77 | Within 10 mV (calibration line) | 1924.768 | 2.16 | 2.16 | Identical | On | On | Same | ¦12.030V 2.160A ¦ | ¦12.030V 2.160A ¦
395 | 1929.6532 | 7788.0801 | 7.78 | 7.79 | Within 10 mV (calibration line) | 1929.6532 | 2.16 | 2.16 | Identical | On | On | Same | ¦12.030V 2.160A ¦ | ¦12.030V 2.160A ¦
396 | 1934.5383 | 7807.7968 | 7.8 | 7.81 | Within 10 mV (calibration line) | 1934.5383 | 2.17 | 2.17 | Identical | On | On | Same | ¦12.030V 2.170A ¦ | ¦12.030V 2.170A ¦
397 | 1939.4235 | 7827.5134 | 7.82 | 7.83 | Within 10 mV (calibration line) | 1939.4235 | 2.17 | 2.18 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 2.170A ¦ | ¦12.030V 2.180A ¦
398 | 1944.3087 | 7847.2301 | 7.84 | 7.85 | Within 10 mV (calibration line) | 1944.3087 | 2.18 | 2.18 | Identical | On | On | Same | ¦12.030V 2.180A ¦ | ¦12.030V 2.180A ¦
399 | 1949.1939 | 7866.9468 | 7.86 | 7.86 | Identical | 1949.1939 | 2.19 | 2.19 | Identical | On | On | Same | ¦12.030V 2.190A ¦ | ¦12.030V 2.190A ¦
400 | 1954.0791 | 7886.6634 | 7.88 | 7.88 | Identical | 1954.0791 | 2.19 | 2.19 | Identical | On | On | Same | ¦12.030V 2.190A ¦ | ¦12.030V 2.190A ¦
401 | 1958.9643 | 7906.3801 | 7.9 | 7.9 | Identical | 1958.9643 | 2.2 | 2.2 | Identical | On | On | Same | ¦12.030V 2.200A ¦ | ¦12.030V 2.200A ¦
402 | 1963.8495 | 7926.0967 | 7.92 | 7.92 | Identical | 1963.8495 | 2.2 | 2.2 | Identical | On | On | Same | ¦12.030V 2.200A ¦ | ¦12.030V 2.200A ¦
403 | 1968.7347 | 7945.8134 | 7.94 | 7.94 | Identical | 1968.7347 | 2.21 | 2.21 | Identical | On | On | Same | ¦12.030V 2.210A ¦ | ¦12.030V 2.210A ¦
404 | 1973.6199 | 7965.53 | 7.96 | 7.96 | Identical | 1973.6199 | 2.21 | 2.21 | Identical | On | On | Same | ¦12.030V 2.210A ¦ | ¦12.030V 2.210A ¦
405 | 1978.5051 | 7985.2467 | 7.98 | 7.98 | Identical | 1978.5051 | 2.22 | 2.22 | Identical | On | On | Same | ¦12.030V 2.220A ¦ | ¦12.030V 2.220A ¦
406 | 1983.3903 | 8004.9634 | 8 | 8 | Identical | 1983.3903 | 2.22 | 2.22 | Identical | On | On | Same | ¦12.030V 2.220A ¦ | ¦12.030V 2.220A ¦
407 | 1988.2755 | 8024.68 | 8.02 | 8.02 | Identical | 1988.2755 | 2.23 | 2.23 | Identical | On | On | Same | ¦12.030V 2.230A ¦ | ¦12.030V 2.230A ¦
408 | 1993.1607 | 8044.3967 | 8.04 | 8.04 | Identical | 1993.1607 | 2.24 | 2.24 | Identical | On | On | Same | ¦12.030V 2.240A ¦ | ¦12.030V 2.240A ¦
409 | 1998.0459 | 8064.1133 | 8.06 | 8.06 | Identical | 1998.0459 | 2.24 | 2.24 | Identical | On | On | Same | ¦12.030V 2.240A ¦ | ¦12.030V 2.240A ¦
410 | 2002.9311 | 8083.83 | 8.08 | 8.08 | Identical | 2002.9311 | 2.25 | 2.25 | Identical | On | On | Same | ¦12.030V 2.250A ¦ | ¦12.030V 2.250A ¦
411 | 2007.8163 | 8103.5467 | 8.1 | 8.1 | Identical | 2007.8163 | 2.25 | 2.25 | Identical | On | On | Same | ¦12.030V 2.250A ¦ | ¦12.030V 2.250A ¦
412 | 2012.7015 | 8123.2633 | 8.12 | 8.12 | Identical | 2012.7015 | 2.26 | 2.26 | Identical | On | On | Same | ¦12.030V 2.260A ¦ | ¦12.030V 2.260A ¦
413 | 2017.5867 | 8142.98 | 8.14 | 8.14 | Identical | 2017.5867 | 2.26 | 2.26 | Identical | On | On | Same | ¦12.030V 2.260A ¦ | ¦12.030V 2.260A ¦
414 | 2022.4719 | 8162.6966 | 8.16 | 8.16 | Identical | 2022.4719 | 2.27 | 2.27 | Identical | On | On | Same | ¦12.030V 2.270A ¦ | ¦12.030V 2.270A ¦
415 | 2027.3571 | 8182.4133 | 8.18 | 8.18 | Identical | 2027.3571 | 2.27 | 2.27 | Identical | On | On | Same | ¦12.030V 2.270A ¦ | ¦12.030V 2.270A ¦
416 | 2032.2423 | 8202.1299 | 8.2 | 8.2 | Identical | 2032.2423 | 2.28 | 2.28 | Identical | On | On | Same | ¦12.030V 2.280A ¦ | ¦12.030V 2.280A ¦
417 | 2037.1275 | 8221.8466 | 8.22 | 8.22 | Identical | 2037.1275 | 2.28 | 2.28 | Identical | On | On | Same | ¦12.030V 2.280A ¦ | ¦12.030V 2.280A ¦
418 | 2042.0127 | 8241.5633 | 8.24 | 8.24 | Identical | 2042.0127 | 2.29 | 2.29 | Identical | On | On | Same | ¦12.030V 2.290A ¦ | ¦12.030V 2.290A ¦
419 | 2046.8979 | 8261.2799 | 8.26 | 8.26 | Identical | 2046.8979 | 2.29 | 2.3 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 2.290A ¦ | ¦12.030V 2.300A ¦
420 | 2051.7831 | 8280.9966 | 8.28 | 8.28 | Identical | 2051.7831 | 2.3 | 2.3 | Identical | On | On | Same | ¦12.030V 2.300A ¦ | ¦12.030V 2.300A ¦
421 | 2056.6683 | 8300.7132 | 8.3 | 8.3 | Identical | 2056.6683 | 2.31 | 2.31 | Identical | On | On | Same | ¦12.030V 2.310A ¦ | ¦12.030V 2.310A ¦
422 | 2061.5535 | 8320.4299 | 8.32 | 8.32 | Identical | 2061.5535 | 2.31 | 2.31 | Identical | On | On | Same | ¦12.030V 2.310A ¦ | ¦12.030V 2.310A ¦
423 | 2066.4387 | 8340.1466 | 8.34 | 8.34 | Identical | 2066.4387 | 2.32 | 2.32 | Identical | On | On | Same | ¦12.030V 2.320A ¦ | ¦12.030V 2.320A ¦
424 | 2071.3239 | 8359.8632 | 8.35 | 8.36 | Within 10 mV (calibration line) | 2071.3239 | 2.32 | 2.32 | Identical | On | On | Same | ¦12.030V 2.320A ¦ | ¦12.030V 2.320A ¦
425 | 2076.2091 | 8379.5799 | 8.37 | 8.38 | Within 10 mV (calibration line) | 2076.2091 | 2.33 | 2.33 | Identical | On | On | Same | ¦12.030V 2.330A ¦ | ¦12.030V 2.330A ¦
426 | 2081.0943 | 8399.2965 | 8.39 | 8.39 | Identical | 2081.0943 | 2.34 | 2.34 | Identical | On | On | Same | ¦12.030V 2.340A ¦ | ¦12.030V 2.340A ¦
427 | 2085.9795 | 8419.0132 | 8.41 | 8.41 | Identical | 2085.9795 | 2.34 | 2.34 | Identical | On | On | Same | ¦12.030V 2.340A ¦ | ¦12.030V 2.340A ¦
428 | 2090.8647 | 8438.7298 | 8.43 | 8.43 | Identical | 2090.8647 | 2.35 | 2.35 | Identical | On | On | Same | ¦12.030V 2.350A ¦ | ¦12.030V 2.350A ¦
429 | 2095.7499 | 8458.4465 | 8.45 | 8.45 | Identical | 2095.7499 | 2.35 | 2.35 | Identical | On | On | Same | ¦12.030V 2.350A ¦ | ¦12.030V 2.350A ¦
430 | 2100.6351 | 8478.1632 | 8.47 | 8.47 | Identical | 2100.6351 | 2.36 | 2.36 | Identical | On | On | Same | ¦12.030V 2.360A ¦ | ¦12.030V 2.360A ¦
431 | 2105.5203 | 8497.8798 | 8.49 | 8.49 | Identical | 2105.5203 | 2.36 | 2.36 | Identical | On | On | Same | ¦12.030V 2.360A ¦ | ¦12.030V 2.360A ¦
432 | 2110.4055 | 8517.5965 | 8.51 | 8.51 | Identical | 2110.4055 | 2.37 | 2.37 | Identical | On | On | Same | ¦12.030V 2.370A ¦ | ¦12.030V 2.370A ¦
433 | 2115.2907 | 8537.3131 | 8.53 | 8.53 | Identical | 2115.2907 | 2.37 | 2.37 | Identical | On | On | Same | ¦12.030V 2.370A ¦ | ¦12.030V 2.370A ¦
434 | 2120.1759 | 8557.0298 | 8.55 | 8.55 | Identical | 2120.1759 | 2.38 | 2.38 | Identical | On | On | Same | ¦12.030V 2.380A ¦ | ¦12.030V 2.380A ¦
435 | 2125.0611 | 8576.7465 | 8.57 | 8.57 | Identical | 2125.0611 | 2.39 | 2.39 | Identical | On | On | Same | ¦12.030V 2.390A ¦ | ¦12.030V 2.390A ¦
436 | 2129.9463 | 8596.4631 | 8.59 | 8.59 | Identical | 2129.9463 | 2.39 | 2.39 | Identical | On | On | Same | ¦12.030V 2.390A ¦ | ¦12.030V 2.390A ¦
437 | 2134.8315 | 8616.1798 | 8.61 | 8.61 | Identical | 2134.8315 | 2.4 | 2.4 | Identical | On | On | Same | ¦12.030V 2.400A ¦ | ¦12.030V 2.400A ¦
438 | 2139.7167 | 8635.8964 | 8.63 | 8.63 | Identical | 2139.7167 | 2.4 | 2.4 | Identical | On | On | Same | ¦12.030V 2.400A ¦ | ¦12.030V 2.400A ¦
439 | 2144.6019 | 8655.6131 | 8.65 | 8.65 | Identical | 2144.6019 | 2.41 | 2.41 | Identical | On | On | Same | ¦12.030V 2.410A ¦ | ¦12.030V 2.410A ¦
440 | 2149.4871 | 8675.3298 | 8.67 | 8.67 | Identical | 2149.4871 | 2.41 | 2.41 | Identical | On | On | Same | ¦12.030V 2.410A ¦ | ¦12.030V 2.410A ¦
441 | 2154.3723 | 8695.0464 | 8.69 | 8.69 | Identical | 2154.3723 | 2.42 | 2.42 | Identical | On | On | Same | ¦12.030V 2.420A ¦ | ¦12.030V 2.420A ¦
442 | 2159.2574 | 8714.7631 | 8.71 | 8.71 | Identical | 2159.2574 | 2.42 | 2.42 | Identical | On | On | Same | ¦12.030V 2.420A ¦ | ¦12.030V 2.420A ¦
443 | 2164.1426 | 8734.4797 | 8.73 | 8.73 | Identical | 2164.1426 | 2.43 | 2.43 | Identical | On | On | Same | ¦12.030V 2.430A ¦ | ¦12.030V 2.430A ¦
444 | 2169.0278 | 8754.1964 | 8.75 | 8.75 | Identical | 2169.0278 | 2.43 | 2.43 | Identical | On | On | Same | ¦12.030V 2.430A ¦ | ¦12.030V 2.430A ¦
445 | 2173.913 | 8773.913 | 8.76 | 8.77 | Within 10 mV (calibration line) | 2173.913 | 2.44 | 2.44 | Identical | On | On | Same | ¦12.030V 2.440A ¦ | ¦12.030V 2.440A ¦
446 | 2178.7982 | 8793.6297 | 8.78 | 8.79 | Within 10 mV (calibration line) | 2178.7982 | 2.44 | 2.45 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 2.440A ¦ | ¦12.030V 2.450A ¦
447 | 2183.6834 | 8813.3464 | 8.8 | 8.81 | Within 10 mV (calibration line) | 2183.6834 | 2.45 | 2.45 | Identical | On | On | Same | ¦12.030V 2.450A ¦ | ¦12.030V 2.450A ¦
448 | 2188.5686 | 8833.063 | 8.82 | 8.83 | Within 10 mV (calibration line) | 2188.5686 | 2.46 | 2.46 | Identical | On | On | Same | ¦12.030V 2.460A ¦ | ¦12.030V 2.460A ¦
449 | 2193.4538 | 8852.7797 | 8.84 | 8.85 | Within 10 mV (calibration line) | 2193.4538 | 2.46 | 2.46 | Identical | On | On | Same | ¦12.030V 2.460A ¦ | ¦12.030V 2.460A ¦
450 | 2198.339 | 8872.4963 | 8.86 | 8.87 | Within 10 mV (calibration line) | 2198.339 | 2.47 | 2.47 | Identical | On | On | Same | ¦12.030V 2.470A ¦ | ¦12.030V 2.470A ¦
451 | 2203.2242 | 8892.213 | 8.88 | 8.89 | Within 10 mV (calibration line) | 2203.2242 | 2.47 | 2.47 | Identical | On | On | Same | ¦12.030V 2.470A ¦ | ¦12.030V 2.470A ¦
452 | 2208.1094 | 8911.9297 | 8.9 | 8.91 | Within 10 mV (calibration line) | 2208.1094 | 2.48 | 2.48 | Identical | On | On | Same | ¦12.030V 2.480A ¦ | ¦12.030V 2.480A ¦
453 | 2212.9946 | 8931.6463 | 8.92 | 8.92 | Identical | 2212.9946 | 2.48 | 2.48 | Identical | On | On | Same | ¦12.030V 2.480A ¦ | ¦12.030V 2.480A ¦
454 | 2217.8798 | 8951.363 | 8.94 | 8.94 | Identical | 2217.8798 | 2.49 | 2.49 | Identical | On | On | Same | ¦12.030V 2.490A ¦ | ¦12.030V 2.490A ¦
455 | 2222.765 | 8971.0796 | 8.96 | 8.96 | Identical | 2222.765 | ERR (999.23) | 2.49 | Gap bug fixed | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V 2.490A ¦
456 | 2227.6502 | 8990.7963 | 8.98 | 8.98 | Identical | 2227.6502 | ERR (999.23) | 2.5 | Gap bug fixed | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V 2.500A ¦
457 | 2232.5354 | 9010.5129 | 9 | 9 | Identical | 2232.5354 | 2.5 | 2.51 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 2.500A0¦ | ¦12.030V 2.510A ¦
458 | 2237.4206 | 9030.2296 | 9.02 | 9.02 | Identical | 2237.4206 | 2.51 | 2.51 | Identical | On | On | Same | ¦12.030V 2.510A0¦ | ¦12.030V 2.510A ¦
459 | 2242.3058 | 9049.9463 | 9.03 | 9.04 | Within 10 mV (calibration line) | 2242.3058 | 2.52 | 2.52 | Identical | On | On | Same | ¦12.030V 2.520A0¦ | ¦12.030V 2.520A ¦
460 | 2247.191 | 9069.6629 | 9.05 | 9.06 | Within 10 mV (calibration line) | 2247.191 | 2.52 | 2.52 | Identical | On | On | Same | ¦12.030V 2.520A0¦ | ¦12.030V 2.520A ¦
461 | 2252.0762 | 9089.3796 | 9.07 | 9.08 | Within 10 mV (calibration line) | 2252.0762 | 2.53 | 2.53 | Identical | On | On | Same | ¦12.030V 2.530A0¦ | ¦12.030V 2.530A ¦
462 | 2256.9614 | 9109.0962 | 9.09 | 9.1 | Within 10 mV (calibration line) | 2256.9614 | 2.53 | 2.53 | Identical | On | On | Same | ¦12.030V 2.530A0¦ | ¦12.030V 2.530A ¦
463 | 2261.8466 | 9128.8129 | 9.11 | 9.12 | Within 10 mV (calibration line) | 2261.8466 | 2.54 | 2.54 | Identical | On | On | Same | ¦12.030V 2.540A0¦ | ¦12.030V 2.540A ¦
464 | 2266.7318 | 9148.5296 | 9.13 | 9.14 | Within 10 mV (calibration line) | 2266.7318 | 2.54 | 2.54 | Identical | On | On | Same | ¦12.030V 2.540A0¦ | ¦12.030V 2.540A ¦
465 | 2271.617 | 9168.2462 | 9.15 | 9.16 | Within 10 mV (calibration line) | 2271.617 | 2.55 | 2.55 | Identical | On | On | Same | ¦12.030V 2.550A0¦ | ¦12.030V 2.550A ¦
466 | 2276.5022 | 9187.9629 | 9.17 | 9.18 | Within 10 mV (calibration line) | 2276.5022 | 2.55 | 2.56 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 2.550A0¦ | ¦12.030V 2.560A ¦
467 | 2281.3874 | 9207.6795 | 9.19 | 9.2 | Within 10 mV (calibration line) | 2281.3874 | 2.56 | 2.56 | Identical | On | On | Same | ¦12.030V 2.560A0¦ | ¦12.030V 2.560A ¦
468 | 2286.2726 | 9227.3962 | 9.21 | 9.22 | Within 10 mV (calibration line) | 2286.2726 | 2.57 | 2.57 | Identical | On | On | Same | ¦12.030V 2.570A0¦ | ¦12.030V 2.570A ¦
469 | 2291.1578 | 9247.1128 | 9.23 | 9.24 | Within 10 mV (calibration line) | 2291.1578 | 2.57 | 2.57 | Identical | On | On | Same | ¦12.030V 2.570A0¦ | ¦12.030V 2.570A ¦
470 | 2296.043 | 9266.8295 | 9.25 | 9.26 | Within 10 mV (calibration line) | 2296.043 | 2.58 | 2.58 | Identical | On | On | Same | ¦12.030V 2.580A0¦ | ¦12.030V 2.580A ¦
471 | 2300.9282 | 9286.5462 | 9.27 | 9.28 | Within 10 mV (calibration line) | 2300.9282 | 2.58 | 2.58 | Identical | On | On | Same | ¦12.030V 2.580A0¦ | ¦12.030V 2.580A ¦
472 | 2305.8134 | 9306.2628 | 9.29 | 9.3 | Within 10 mV (calibration line) | 2305.8134 | 2.59 | 2.59 | Identical | On | On | Same | ¦12.030V 2.590A0¦ | ¦12.030V 2.590A ¦
473 | 2310.6986 | 9325.9795 | 9.31 | 9.32 | Within 10 mV (calibration line) | 2310.6986 | 2.59 | 2.59 | Identical | On | On | Same | ¦12.030V 2.590A0¦ | ¦12.030V 2.590A ¦
474 | 2315.5838 | 9345.6961 | 9.33 | 9.34 | Within 10 mV (calibration line) | 2315.5838 | 2.6 | 2.6 | Identical | On | On | Same | ¦12.030V 2.600A0¦ | ¦12.030V 2.600A ¦
475 | 2320.469 | 9365.4128 | 9.35 | 9.36 | Within 10 mV (calibration line) | 2320.469 | 2.6 | 2.6 | Identical | On | On | Same | ¦12.030V 2.600A0¦ | ¦12.030V 2.600A ¦
476 | 2325.3542 | 9385.1295 | 9.37 | 9.38 | Within 10 mV (calibration line) | 2325.3542 | 2.61 | 2.61 | Identical | On | On | Same | ¦12.030V 2.610A0¦ | ¦12.030V 2.610A ¦
477 | 2330.2394 | 9404.8461 | 9.39 | 9.4 | Within 10 mV (calibration line) | 2330.2394 | 2.61 | 2.61 | Identical | On | On | Same | ¦12.030V 2.610A0¦ | ¦12.030V 2.610A ¦
478 | 2335.1246 | 9424.5628 | 9.41 | 9.42 | Within 10 mV (calibration line) | 2335.1246 | 2.62 | 2.62 | Identical | On | On | Same | ¦12.030V 2.620A0¦ | ¦12.030V 2.620A ¦
479 | 2340.0098 | 9444.2794 | 9.43 | 9.44 | Within 10 mV (calibration line) | 2340.0098 | 2.63 | 2.63 | Identical | On | On | Same | ¦12.030V 2.630A0¦ | ¦12.030V 2.630A ¦
480 | 2344.895 | 9463.9961 | 9.45 | 9.46 | Within 10 mV (calibration line) | 2344.895 | 2.63 | 2.63 | Identical | On | On | Same | ¦12.030V 2.630A0¦ | ¦12.030V 2.630A ¦
481 | 2349.7802 | 9483.7128 | 9.47 | 9.47 | Identical | 2349.7802 | 2.64 | 2.64 | Identical | On | On | Same | ¦12.030V 2.640A0¦ | ¦12.030V 2.640A ¦
482 | 2354.6654 | 9503.4294 | 9.49 | 9.49 | Identical | 2354.6654 | 2.64 | 2.64 | Identical | On | On | Same | ¦12.030V 2.640A0¦ | ¦12.030V 2.640A ¦
483 | 2359.5506 | 9523.1461 | 9.51 | 9.51 | Identical | 2359.5506 | 2.65 | 2.65 | Identical | On | On | Same | ¦12.030V 2.650A0¦ | ¦12.030V 2.650A ¦
484 | 2364.4358 | 9542.8627 | 9.53 | 9.53 | Identical | 2364.4358 | 2.65 | 2.65 | Identical | On | On | Same | ¦12.030V 2.650A0¦ | ¦12.030V 2.650A ¦
485 | 2369.321 | 9562.5794 | 9.55 | 9.55 | Identical | 2369.321 | 2.66 | 2.66 | Identical | On | On | Same | ¦12.030V 2.660A0¦ | ¦12.030V 2.660A ¦
486 | 2374.2062 | 9582.296 | 9.57 | 9.57 | Identical | 2374.2062 | 2.66 | 2.66 | Identical | On | On | Same | ¦12.030V 2.660A0¦ | ¦12.030V 2.660A ¦
487 | 2379.0914 | 9602.0127 | 9.59 | 9.59 | Identical | 2379.0914 | 2.67 | 2.67 | Identical | On | On | Same | ¦12.030V 2.670A0¦ | ¦12.030V 2.670A ¦
488 | 2383.9766 | 9621.7294 | 9.61 | 9.61 | Identical | 2383.9766 | 2.67 | 2.67 | Identical | On | On | Same | ¦12.030V 2.670A0¦ | ¦12.030V 2.670A ¦
489 | 2388.8617 | 9641.446 | 9.63 | 9.63 | Identical | 2388.8617 | 2.68 | 2.68 | Identical | On | On | Same | ¦12.030V 2.680A0¦ | ¦12.030V 2.680A ¦
490 | 2393.7469 | 9661.1627 | 9.65 | 9.65 | Identical | 2393.7469 | 2.68 | 2.69 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 2.680A0¦ | ¦12.030V 2.690A ¦
491 | 2398.6321 | 9680.8793 | 9.67 | 9.67 | Identical | 2398.6321 | 2.69 | 2.69 | Identical | On | On | Same | ¦12.030V 2.690A0¦ | ¦12.030V 2.690A ¦
492 | 2403.5173 | 9700.596 | 9.69 | 9.69 | Identical | 2403.5173 | 2.7 | 2.7 | Identical | On | On | Same | ¦12.030V 2.700A0¦ | ¦12.030V 2.700A ¦
493 | 2408.4025 | 9720.3127 | 9.71 | 9.71 | Identical | 2408.4025 | 2.7 | 2.7 | Identical | On | On | Same | ¦12.030V 2.700A0¦ | ¦12.030V 2.700A ¦
494 | 2413.2877 | 9740.0293 | 9.73 | 9.73 | Identical | 2413.2877 | 2.71 | 2.71 | Identical | On | On | Same | ¦12.030V 2.710A0¦ | ¦12.030V 2.710A ¦
495 | 2418.1729 | 9759.746 | 9.74 | 9.75 | Within 10 mV (calibration line) | 2418.1729 | 2.71 | 2.71 | Identical | On | On | Same | ¦12.030V 2.710A0¦ | ¦12.030V 2.710A ¦
496 | 2423.0581 | 9779.4626 | 9.76 | 9.77 | Within 10 mV (calibration line) | 2423.0581 | 2.72 | 2.72 | Identical | On | On | Same | ¦12.030V 2.720A0¦ | ¦12.030V 2.720A ¦
497 | 2427.9433 | 9799.1793 | 9.78 | 9.79 | Within 10 mV (calibration line) | 2427.9433 | 2.72 | 2.72 | Identical | On | On | Same | ¦12.030V 2.720A0¦ | ¦12.030V 2.720A ¦
498 | 2432.8285 | 9818.8959 | 9.8 | 9.81 | Within 10 mV (calibration line) | 2432.8285 | 2.73 | 2.73 | Identical | On | On | Same | ¦12.030V 2.730A0¦ | ¦12.030V 2.730A ¦
499 | 2437.7137 | 9838.6126 | 9.82 | 9.83 | Within 10 mV (calibration line) | 2437.7137 | 2.73 | 2.74 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 2.730A0¦ | ¦12.030V 2.740A ¦
500 | 2442.5989 | 9858.3293 | 9.84 | 9.85 | Within 10 mV (calibration line) | 2442.5989 | 2.74 | 2.74 | Identical | On | On | Same | ¦12.030V 2.740A0¦ | ¦12.030V 2.740A ¦
501 | 2447.4841 | 9878.0459 | 9.86 | 9.87 | Within 10 mV (calibration line) | 2447.4841 | 2.75 | 2.75 | Identical | On | On | Same | ¦12.030V 2.750A0¦ | ¦12.030V 2.750A ¦
502 | 2452.3693 | 9897.7626 | 9.88 | 9.89 | Within 10 mV (calibration line) | 2452.3693 | 2.75 | 2.75 | Identical | On | On | Same | ¦12.030V 2.750A0¦ | ¦12.030V 2.750A ¦
503 | 2457.2545 | 9917.4792 | 9.9 | 9.91 | Within 10 mV (calibration line) | 2457.2545 | 2.76 | 2.76 | Identical | On | On | Same | ¦12.030V 2.760A0¦ | ¦12.030V 2.760A ¦
504 | 2462.1397 | 9937.1959 | 9.92 | 9.93 | Within 10 mV (calibration line) | 2462.1397 | 2.76 | 2.76 | Identical | On | On | Same | ¦12.030V 2.760A0¦ | ¦12.030V 2.760A ¦
505 | 2467.0249 | 9956.9126 | 9.94 | 9.95 | Within 10 mV (calibration line) | 2467.0249 | 2.77 | 2.77 | Identical | On | On | Same | ¦12.030V 2.770A0¦ | ¦12.030V 2.770A ¦
506 | 2471.9101 | 9976.6292 | 9.96 | 9.97 | Within 10 mV (calibration line) | 2471.9101 | 2.77 | 2.77 | Identical | On | On | Same | ¦12.030V 2.770A0¦ | ¦12.030V 2.770A ¦
507 | 2476.7953 | 9996.3459 | 9.98 | 9.99 | Within 10 mV (calibration line) | 2476.7953 | 2.78 | 2.78 | Identical | On | On | Same | ¦12.030V 2.780A0¦ | ¦12.030V 2.780A ¦
508 | 2481.6805 | 10016.0625 | 10 | 10 | Identical | 2481.6805 | 2.78 | 2.78 | Identical | On | On | Same | ¦12.030V 2.780A0¦ | ¦12.030V 2.780A ¦
509 | 2486.5657 | 10035.7792 | 10.02 | 10.02 | Identical | 2486.5657 | 2.79 | 2.79 | Identical | On | On | Same | ¦12.030V 2.790A0¦ | ¦12.030V 2.790A ¦
510 | 2491.4509 | 10055.4958 | 10.04 | 10.04 | Identical | 2491.4509 | 2.79 | 2.79 | Identical | On | On | Same | ¦12.030V 2.790A0¦ | ¦12.030V 2.790A ¦
511 | 2496.3361 | 10075.2125 | 10.06 | 10.06 | Identical | 2496.3361 | 2.8 | 2.8 | Identical | On | On | Same | ¦12.030V 2.800A0¦ | ¦12.030V 2.800A ¦
512 | 2501.2213 | 10094.9292 | 10.08 | 10.08 | Identical | 2501.2213 | 2.81 | 2.81 | Identical | On | On | Same | ¦12.030V 2.810A0¦ | ¦12.030V 2.810A ¦
513 | 2506.1065 | 10114.6458 | 10.1 | 10.1 | Identical | 2506.1065 | 2.81 | 2.81 | Identical | On | On | Same | ¦12.030V 2.810A0¦ | ¦12.030V 2.810A ¦
514 | 2510.9917 | 10134.3625 | 10.12 | 10.12 | Identical | 2510.9917 | 2.82 | 2.82 | Identical | On | On | Same | ¦12.030V 2.820A0¦ | ¦12.030V 2.820A ¦
515 | 2515.8769 | 10154.0791 | 10.14 | 10.14 | Identical | 2515.8769 | 2.82 | 2.82 | Identical | On | On | Same | ¦12.030V 2.820A0¦ | ¦12.030V 2.820A ¦
516 | 2520.7621 | 10173.7958 | 10.16 | 10.16 | Identical | 2520.7621 | 2.83 | 2.83 | Identical | On | On | Same | ¦12.030V 2.830A0¦ | ¦12.030V 2.830A ¦
517 | 2525.6473 | 10193.5125 | 10.18 | 10.18 | Identical | 2525.6473 | 2.83 | 2.83 | Identical | On | On | Same | ¦12.030V 2.830A0¦ | ¦12.030V 2.830A ¦
518 | 2530.5325 | 10213.2291 | 10.2 | 10.2 | Identical | 2530.5325 | 2.84 | 2.84 | Identical | On | On | Same | ¦12.030V 2.840A0¦ | ¦12.030V 2.840A ¦
519 | 2535.4177 | 10232.9458 | 10.22 | 10.22 | Identical | 2535.4177 | 2.84 | 2.85 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 2.840A0¦ | ¦12.030V 2.850A ¦
520 | 2540.3029 | 10252.6624 | 10.24 | 10.24 | Identical | 2540.3029 | 2.85 | 2.85 | Identical | On | On | Same | ¦12.030V 2.850A0¦ | ¦12.030V 2.850A ¦
521 | 2545.1881 | 10272.3791 | 10.26 | 10.26 | Identical | 2545.1881 | 2.86 | 2.86 | Identical | On | On | Same | ¦12.030V 2.860A0¦ | ¦12.030V 2.860A ¦
522 | 2550.0733 | 10292.0957 | 10.28 | 10.28 | Identical | 2550.0733 | 2.86 | 2.86 | Identical | On | On | Same | ¦12.030V 2.860A0¦ | ¦12.030V 2.860A ¦
523 | 2554.9585 | 10311.8124 | 10.3 | 10.3 | Identical | 2554.9585 | 2.87 | 2.87 | Identical | On | On | Same | ¦12.030V 2.870A0¦ | ¦12.030V 2.870A ¦
524 | 2559.8437 | 10331.5291 | 10.32 | 10.32 | Identical | 2559.8437 | 2.87 | 2.87 | Identical | On | On | Same | ¦12.030V 2.870A0¦ | ¦12.030V 2.870A ¦
525 | 2564.7289 | 10351.2457 | 10.34 | 10.34 | Identical | 2564.7289 | 2.88 | 2.88 | Identical | On | On | Same | ¦12.030V 2.880A0¦ | ¦12.030V 2.880A ¦
526 | 2569.6141 | 10370.9624 | 10.36 | 10.36 | Identical | 2569.6141 | 2.88 | 2.89 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 2.880A0¦ | ¦12.030V 2.890A ¦
527 | 2574.4993 | 10390.679 | 10.38 | 10.38 | Identical | 2574.4993 | 2.89 | 2.89 | Identical | On | On | Same | ¦12.030V 2.890A0¦ | ¦12.030V 2.890A ¦
528 | 2579.3845 | 10410.3957 | 10.4 | 10.4 | Identical | 2579.3845 | 2.9 | 2.9 | Identical | On | On | Same | ¦12.030V 2.900A0¦ | ¦12.030V 2.900A ¦
529 | 2584.2697 | 10430.1124 | 10.42 | 10.42 | Identical | 2584.2697 | 2.9 | 2.9 | Identical | On | On | Same | ¦12.030V 2.900A0¦ | ¦12.030V 2.900A ¦
530 | 2589.1549 | 10449.829 | 10.43 | 10.44 | Within 10 mV (calibration line) | 2589.1549 | 2.91 | 2.91 | Identical | On | On | Same | ¦12.030V 2.910A0¦ | ¦12.030V 2.910A ¦
531 | 2594.0401 | 10469.5457 | 10.45 | 10.46 | Within 10 mV (calibration line) | 2594.0401 | 2.91 | 2.91 | Identical | On | On | Same | ¦12.030V 2.910A0¦ | ¦12.030V 2.910A ¦
532 | 2598.9253 | 10489.2623 | 10.47 | 10.48 | Within 10 mV (calibration line) | 2598.9253 | 2.92 | 2.92 | Identical | On | On | Same | ¦12.030V 2.920A0¦ | ¦12.030V 2.920A ¦
533 | 2603.8105 | 10508.979 | 10.49 | 10.5 | Within 10 mV (calibration line) | 2603.8105 | 2.92 | 2.92 | Identical | On | On | Same | ¦12.030V 2.920A0¦ | ¦12.030V 2.920A ¦
534 | 2608.6957 | 10528.6957 | 10.51 | 10.52 | Within 10 mV (calibration line) | 2608.6957 | 2.93 | 2.93 | Identical | On | On | Same | ¦12.030V 2.930A0¦ | ¦12.030V 2.930A ¦
535 | 2613.5809 | 10548.4123 | 10.53 | 10.53 | Identical | 2613.5809 | 2.93 | 2.94 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 2.930A0¦ | ¦12.030V 2.940A ¦
536 | 2618.466 | 10568.129 | 10.55 | 10.55 | Identical | 2618.466 | 2.94 | 2.94 | Identical | On | On | Same | ¦12.030V 2.940A0¦ | ¦12.030V 2.940A ¦
537 | 2623.3512 | 10587.8456 | 10.57 | 10.57 | Identical | 2623.3512 | 2.95 | 2.95 | Identical | On | On | Same | ¦12.030V 2.950A0¦ | ¦12.030V 2.950A ¦
538 | 2628.2364 | 10607.5623 | 10.59 | 10.59 | Identical | 2628.2364 | 2.95 | 2.95 | Identical | On | On | Same | ¦12.030V 2.950A0¦ | ¦12.030V 2.950A ¦
539 | 2633.1216 | 10627.2789 | 10.61 | 10.61 | Identical | 2633.1216 | 2.96 | 2.96 | Identical | On | On | Same | ¦12.030V 2.960A0¦ | ¦12.030V 2.960A ¦
540 | 2638.0068 | 10646.9956 | 10.63 | 10.63 | Identical | 2638.0068 | 2.96 | 2.96 | Identical | On | On | Same | ¦12.030V 2.960A0¦ | ¦12.030V 2.960A ¦
541 | 2642.892 | 10666.7123 | 10.65 | 10.65 | Identical | 2642.892 | 2.97 | 2.97 | Identical | On | On | Same | ¦12.030V 2.970A0¦ | ¦12.030V 2.970A ¦
542 | 2647.7772 | 10686.4289 | 10.67 | 10.67 | Identical | 2647.7772 | 2.97 | 2.97 | Identical | On | On | Same | ¦12.030V 2.970A0¦ | ¦12.030V 2.970A ¦
543 | 2652.6624 | 10706.1456 | 10.69 | 10.69 | Identical | 2652.6624 | 2.98 | 2.98 | Identical | On | On | Same | ¦12.030V 2.980A0¦ | ¦12.030V 2.980A ¦
544 | 2657.5476 | 10725.8622 | 10.71 | 10.71 | Identical | 2657.5476 | 2.98 | 2.98 | Identical | On | On | Same | ¦12.030V 2.980A0¦ | ¦12.030V 2.980A ¦
545 | 2662.4328 | 10745.5789 | 10.73 | 10.73 | Identical | 2662.4328 | 2.99 | 2.99 | Identical | On | On | Same | ¦12.030V 2.990A0¦ | ¦12.030V 2.990A ¦
546 | 2667.318 | 10765.2956 | 10.75 | 10.75 | Identical | 2667.318 | 2.99 | 2.99 | Identical | On | On | Same | ¦12.030V 2.990A0¦ | ¦12.030V 2.990A ¦
547 | 2672.2032 | 10785.0122 | 10.77 | 10.77 | Identical | 2672.2032 | 3 | 3 | Identical | On | On | Same | ¦12.030V 3.000A0¦ | ¦12.030V 3.000A ¦
548 | 2677.0884 | 10804.7289 | 10.79 | 10.79 | Identical | 2677.0884 | 3.01 | 3.01 | Identical | On | On | Same | ¦12.030V 3.010A0¦ | ¦12.030V 3.010A ¦
549 | 2681.9736 | 10824.4455 | 10.81 | 10.81 | Identical | 2681.9736 | 3.01 | 3.01 | Identical | On | On | Same | ¦12.030V 3.010A0¦ | ¦12.030V 3.010A ¦
550 | 2686.8588 | 10844.1622 | 10.83 | 10.83 | Identical | 2686.8588 | 3.02 | 3.02 | Identical | On | On | Same | ¦12.030V 3.020A0¦ | ¦12.030V 3.020A ¦
551 | 2691.744 | 10863.8788 | 10.85 | 10.85 | Identical | 2691.744 | 3.02 | 3.02 | Identical | On | On | Same | ¦12.030V 3.020A0¦ | ¦12.030V 3.020A ¦
552 | 2696.6292 | 10883.5955 | 10.87 | 10.87 | Identical | 2696.6292 | 3.03 | 3.03 | Identical | On | On | Same | ¦12.030V 3.030A0¦ | ¦12.030V 3.030A ¦
553 | 2701.5144 | 10903.3122 | 10.89 | 10.89 | Identical | 2701.5144 | 3.03 | 3.03 | Identical | On | On | Same | ¦12.030V 3.030A0¦ | ¦12.030V 3.030A ¦
554 | 2706.3996 | 10923.0288 | 10.91 | 10.91 | Identical | 2706.3996 | 3.04 | 3.04 | Identical | On | On | Same | ¦12.030V 3.040A0¦ | ¦12.030V 3.040A ¦
555 | 2711.2848 | 10942.7455 | 10.93 | 10.93 | Identical | 2711.2848 | 3.04 | 3.04 | Identical | On | On | Same | ¦12.030V 3.040A0¦ | ¦12.030V 3.040A ¦
556 | 2716.17 | 10962.4621 | 10.95 | 10.95 | Identical | 2716.17 | 3.05 | 3.05 | Identical | On | On | Same | ¦12.030V 3.050A0¦ | ¦12.030V 3.050A ¦
557 | 2721.0552 | 10982.1788 | 10.97 | 10.97 | Identical | 2721.0552 | 3.05 | 3.05 | Identical | On | On | Same | ¦12.030V 3.050A0¦ | ¦12.030V 3.050A ¦
558 | 2725.9404 | 11001.8955 | 10.99 | 10.99 | Identical | 2725.9404 | 3.06 | 3.06 | Identical | On | On | Same | ¦12.030V 3.060A0¦ | ¦12.030V 3.060A ¦
559 | 2730.8256 | 11021.6121 | 11.01 | 11.01 | Identical | 2730.8256 | 3.06 | 3.07 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 3.060A0¦ | ¦12.030V 3.070A ¦
560 | 2735.7108 | 11041.3288 | 11.03 | 11.03 | Identical | 2735.7108 | 3.07 | 3.07 | Identical | On | On | Same | ¦12.030V 3.070A0¦ | ¦12.030V 3.070A ¦
561 | 2740.596 | 11061.0454 | 11.05 | 11.05 | Identical | 2740.596 | 3.08 | 3.08 | Identical | On | On | Same | ¦12.030V 3.080A0¦ | ¦12.030V 3.080A ¦
562 | 2745.4812 | 11080.7621 | 11.07 | 11.07 | Identical | 2745.4812 | 3.08 | 3.08 | Identical | On | On | Same | ¦12.030V 3.080A0¦ | ¦12.030V 3.080A ¦
563 | 2750.3664 | 11100.4787 | 11.09 | 11.08 | Within 10 mV (calibration line) | 2750.3664 | 3.09 | 3.09 | Identical | On | On | Same | ¦12.030V 3.090A0¦ | ¦12.030V 3.090A ¦
564 | 2755.2516 | 11120.1954 | 11.11 | 11.1 | Within 10 mV (calibration line) | 2755.2516 | 3.09 | 3.09 | Identical | On | On | Same | ¦12.030V 3.090A0¦ | ¦12.030V 3.090A ¦
565 | 2760.1368 | 11139.9121 | 11.12 | 11.12 | Identical | 2760.1368 | 3.1 | 3.1 | Identical | On | On | Same | ¦12.030V 3.100A0¦ | ¦12.030V 3.100A ¦
566 | 2765.022 | 11159.6287 | 11.14 | 11.14 | Identical | 2765.022 | 3.1 | 3.1 | Identical | On | On | Same | ¦12.030V 3.100A0¦ | ¦12.030V 3.100A ¦
567 | 2769.9072 | 11179.3454 | 11.16 | 11.16 | Identical | 2769.9072 | 3.11 | 3.11 | Identical | On | On | Same | ¦12.030V 3.110A0¦ | ¦12.030V 3.110A ¦
568 | 2774.7924 | 11199.062 | 11.18 | 11.18 | Identical | 2774.7924 | 3.11 | 3.12 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 3.110A0¦ | ¦12.030V 3.120A ¦
569 | 2779.6776 | 11218.7787 | 11.2 | 11.2 | Identical | 2779.6776 | 3.12 | 3.12 | Identical | On | On | Same | ¦12.030V 3.120A0¦ | ¦12.030V 3.120A ¦
570 | 2784.5628 | 11238.4954 | 11.22 | 11.22 | Identical | 2784.5628 | 3.13 | 3.13 | Identical | On | On | Same | ¦12.030V 3.130A0¦ | ¦12.030V 3.130A ¦
571 | 2789.448 | 11258.212 | 11.24 | 11.24 | Identical | 2789.448 | 3.13 | 3.13 | Identical | On | On | Same | ¦12.030V 3.130A0¦ | ¦12.030V 3.130A ¦
572 | 2794.3332 | 11277.9287 | 11.26 | 11.26 | Identical | 2794.3332 | 3.14 | 3.14 | Identical | On | On | Same | ¦12.030V 3.140A0¦ | ¦12.030V 3.140A ¦
573 | 2799.2184 | 11297.6453 | 11.28 | 11.28 | Identical | 2799.2184 | 3.14 | 3.14 | Identical | On | On | Same | ¦12.030V 3.140A0¦ | ¦12.030V 3.140A ¦
574 | 2804.1036 | 11317.362 | 11.3 | 11.3 | Identical | 2804.1036 | 3.15 | 3.15 | Identical | On | On | Same | ¦12.030V 3.150A0¦ | ¦12.030V 3.150A ¦
575 | 2808.9888 | 11337.0787 | 11.32 | 11.32 | Identical | 2808.9888 | 3.15 | 3.15 | Identical | On | On | Same | ¦12.030V 3.150A0¦ | ¦12.030V 3.150A ¦
576 | 2813.874 | 11356.7953 | 11.34 | 11.34 | Identical | 2813.874 | 3.16 | 3.16 | Identical | On | On | Same | ¦12.030V 3.160A0¦ | ¦12.030V 3.160A ¦
577 | 2818.7592 | 11376.512 | 11.36 | 11.36 | Identical | 2818.7592 | 3.16 | 3.16 | Identical | On | On | Same | ¦12.030V 3.160A0¦ | ¦12.030V 3.160A ¦
578 | 2823.6444 | 11396.2286 | 11.38 | 11.38 | Identical | 2823.6444 | 3.17 | 3.17 | Identical | On | On | Same | ¦12.030V 3.170A0¦ | ¦12.030V 3.170A ¦
579 | 2828.5296 | 11415.9453 | 11.4 | 11.4 | Identical | 2828.5296 | 3.17 | 3.18 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 3.170A0¦ | ¦12.030V 3.180A ¦
580 | 2833.4148 | 11435.6619 | 11.42 | 11.42 | Identical | 2833.4148 | 3.18 | 3.18 | Identical | On | On | Same | ¦12.030V 3.180A0¦ | ¦12.030V 3.180A ¦
581 | 2838.3 | 11455.3786 | 11.44 | 11.44 | Identical | 2838.3 | 3.19 | 3.19 | Identical | On | On | Same | ¦12.030V 3.190A0¦ | ¦12.030V 3.190A ¦
582 | 2843.1851 | 11475.0953 | 11.46 | 11.46 | Identical | 2843.1851 | 3.19 | 3.19 | Identical | On | On | Same | ¦12.030V 3.190A0¦ | ¦12.030V 3.190A ¦
583 | 2848.0703 | 11494.8119 | 11.48 | 11.48 | Identical | 2848.0703 | 3.2 | 3.2 | Identical | On | On | Same | ¦12.030V 3.200A0¦ | ¦12.030V 3.200A ¦
584 | 2852.9555 | 11514.5286 | 11.5 | 11.5 | Identical | 2852.9555 | 3.2 | 3.2 | Identical | On | On | Same | ¦12.030V 3.200A0¦ | ¦12.030V 3.200A ¦
585 | 2857.8407 | 11534.2452 | 11.52 | 11.52 | Identical | 2857.8407 | 3.21 | 3.21 | Identical | On | On | Same | ¦12.030V 3.210A0¦ | ¦12.030V 3.210A ¦
586 | 2862.7259 | 11553.9619 | 11.54 | 11.54 | Identical | 2862.7259 | 3.21 | 3.21 | Identical | On | On | Same | ¦12.030V 3.210A0¦ | ¦12.030V 3.210A ¦
587 | 2867.6111 | 11573.6786 | 11.56 | 11.56 | Identical | 2867.6111 | 3.22 | 3.22 | Identical | On | On | Same | ¦12.030V 3.220A0¦ | ¦12.030V 3.220A ¦
588 | 2872.4963 | 11593.3952 | 11.58 | 11.58 | Identical | 2872.4963 | 3.22 | 3.22 | Identical | On | On | Same | ¦12.030V 3.220A0¦ | ¦12.030V 3.220A ¦
589 | 2877.3815 | 11613.1119 | 11.6 | 11.6 | Identical | 2877.3815 | 3.23 | 3.23 | Identical | On | On | Same | ¦12.030V 3.230A0¦ | ¦12.030V 3.230A ¦
590 | 2882.2667 | 11632.8285 | 11.62 | 11.61 | Within 10 mV (calibration line) | 2882.2667 | 3.23 | 3.23 | Identical | On | On | Same | ¦12.030V 3.230A0¦ | ¦12.030V 3.230A ¦
591 | 2887.1519 | 11652.5452 | 11.64 | 11.63 | Within 10 mV (calibration line) | 2887.1519 | 3.24 | 3.24 | Identical | On | On | Same | ¦12.030V 3.240A0¦ | ¦12.030V 3.240A ¦
592 | 2892.0371 | 11672.2618 | 11.66 | 11.65 | Within 10 mV (calibration line) | 2892.0371 | 3.25 | 3.25 | Identical | On | On | Same | ¦12.030V 3.250A0¦ | ¦12.030V 3.250A ¦
593 | 2896.9223 | 11691.9785 | 11.68 | 11.67 | Within 10 mV (calibration line) | 2896.9223 | 3.25 | 3.25 | Identical | On | On | Same | ¦12.030V 3.250A0¦ | ¦12.030V 3.250A ¦
594 | 2901.8075 | 11711.6952 | 11.7 | 11.69 | Within 10 mV (calibration line) | 2901.8075 | 3.26 | 3.26 | Identical | On | On | Same | ¦12.030V 3.260A0¦ | ¦12.030V 3.260A ¦
595 | 2906.6927 | 11731.4118 | 11.72 | 11.71 | Within 10 mV (calibration line) | 2906.6927 | 3.26 | 3.26 | Identical | On | On | Same | ¦12.030V 3.260A0¦ | ¦12.030V 3.260A ¦
596 | 2911.5779 | 11751.1285 | 11.74 | 11.73 | Within 10 mV (calibration line) | 2911.5779 | 3.27 | 3.27 | Identical | On | On | Same | ¦12.030V 3.270A0¦ | ¦12.030V 3.270A ¦
597 | 2916.4631 | 11770.8451 | 11.76 | 11.75 | Within 10 mV (calibration line) | 2916.4631 | 3.28 | 3.28 | Identical | On | On | Same | ¦12.030V 3.280A0¦ | ¦12.030V 3.280A ¦
598 | 2921.3483 | 11790.5618 | 11.78 | 11.77 | Within 10 mV (calibration line) | 2921.3483 | 3.28 | 3.28 | Identical | On | On | Same | ¦12.030V 3.280A0¦ | ¦12.030V 3.280A ¦
599 | 2926.2335 | 11810.2785 | 11.8 | 11.79 | Within 10 mV (calibration line) | 2926.2335 | 3.29 | 3.29 | Identical | On | On | Same | ¦12.030V 3.290A0¦ | ¦12.030V 3.290A ¦
600 | 2931.1187 | 11829.9951 | 11.81 | 11.81 | Identical | 2931.1187 | 3.29 | 3.29 | Identical | On | On | Same | ¦12.030V 3.290A0¦ | ¦12.030V 3.290A ¦
601 | 2936.0039 | 11849.7118 | 11.83 | 11.83 | Identical | 2936.0039 | 3.3 | 3.3 | Identical | On | On | Same | ¦12.030V 3.300A0¦ | ¦12.030V 3.300A ¦
602 | 2940.8891 | 11869.4284 | 11.85 | 11.85 | Identical | 2940.8891 | 3.3 | 3.3 | Identical | On | On | Same | ¦12.030V 3.300A0¦ | ¦12.030V 3.300A ¦
603 | 2945.7743 | 11889.1451 | 11.87 | 11.87 | Identical | 2945.7743 | 3.31 | 3.31 | Identical | On | On | Same | ¦12.030V 3.310A0¦ | ¦12.030V 3.310A ¦
604 | 2950.6595 | 11908.8617 | 11.89 | 11.89 | Identical | 2950.6595 | 3.31 | 3.32 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 3.310A0¦ | ¦12.030V 3.320A ¦
605 | 2955.5447 | 11928.5784 | 11.91 | 11.91 | Identical | 2955.5447 | 3.32 | 3.32 | Identical | On | On | Same | ¦12.030V 3.320A0¦ | ¦12.030V 3.320A ¦
606 | 2960.4299 | 11948.2951 | 11.93 | 11.93 | Identical | 2960.4299 | 3.33 | 3.33 | Identical | On | On | Same | ¦12.030V 3.330A0¦ | ¦12.030V 3.330A ¦
607 | 2965.3151 | 11968.0117 | 11.95 | 11.95 | Identical | 2965.3151 | 3.33 | 3.33 | Identical | On | On | Same | ¦12.030V 3.330A0¦ | ¦12.030V 3.330A ¦
608 | 2970.2003 | 11987.7284 | 11.97 | 11.97 | Identical | 2970.2003 | 3.34 | 3.34 | Identical | On | On | Same | ¦12.030V 3.340A0¦ | ¦12.030V 3.340A ¦
609 | 2975.0855 | 12007.445 | 11.99 | 11.99 | Identical | 2975.0855 | 3.34 | 3.34 | Identical | On | On | Same | ¦12.030V 3.340A0¦ | ¦12.030V 3.340A ¦
610 | 2979.9707 | 12027.1617 | 12.01 | 12.01 | Identical | 2979.9707 | 3.35 | 3.35 | Identical | On | On | Same | ¦12.030V 3.350A0¦ | ¦12.030V 3.350A ¦
611 | 2984.8559 | 12046.8784 | 12.03 | 12.03 | Identical | 2984.8559 | 3.35 | 3.35 | Identical | On | On | Same | ¦12.030V 3.350A0¦ | ¦12.030V 3.350A ¦
612 | 2989.7411 | 12066.595 | 12.05 | 12.05 | Identical | 2989.7411 | 3.36 | 3.36 | Identical | On | On | Same | ¦12.030V 3.360A0¦ | ¦12.030V 3.360A ¦
613 | 2994.6263 | 12086.3117 | 12.07 | 12.07 | Identical | 2994.6263 | 3.36 | 3.36 | Identical | On | On | Same | ¦12.030V 3.360A0¦ | ¦12.030V 3.360A ¦
614 | 2999.5115 | 12106.0283 | 12.09 | 12.09 | Identical | 2999.5115 | 3.37 | 3.37 | Identical | On | On | Same | ¦12.030V 3.370A0¦ | ¦12.030V 3.370A ¦
615 | 3004.3967 | 12125.745 | 12.11 | 12.11 | Identical | 3004.3967 | 3.37 | 3.37 | Identical | On | On | Same | ¦12.030V 3.370A0¦ | ¦12.030V 3.370A ¦
616 | 3009.2819 | 12145.4617 | 12.13 | 12.13 | Identical | 3009.2819 | 3.38 | 3.38 | Identical | On | On | Same | ¦12.030V 3.380A0¦ | ¦12.030V 3.380A ¦
617 | 3014.1671 | 12165.1783 | 12.15 | 12.14 | Within 10 mV (calibration line) | 3014.1671 | 3.39 | 3.39 | Identical | On | On | Same | ¦12.030V 3.390A0¦ | ¦12.030V 3.390A ¦
618 | 3019.0523 | 12184.895 | 12.17 | 12.16 | Within 10 mV (calibration line) | 3019.0523 | 3.39 | 3.39 | Identical | On | On | Same | ¦12.030V 3.390A0¦ | ¦12.030V 3.390A ¦
619 | 3023.9375 | 12204.6116 | 12.19 | 12.18 | Within 10 mV (calibration line) | 3023.9375 | 3.4 | 3.4 | Identical | On | On | Same | ¦12.030V 3.400A0¦ | ¦12.030V 3.400A ¦
620 | 3028.8227 | 12224.3283 | 12.21 | 12.2 | Within 10 mV (calibration line) | 3028.8227 | 3.4 | 3.4 | Identical | On | On | Same | ¦12.030V 3.400A0¦ | ¦12.030V 3.400A ¦
621 | 3033.7079 | 12244.0449 | 12.23 | 12.22 | Within 10 mV (calibration line) | 3033.7079 | 3.41 | 3.41 | Identical | On | On | Same | ¦12.030V 3.410A0¦ | ¦12.030V 3.410A ¦
622 | 3038.5931 | 12263.7616 | 12.25 | 12.24 | Within 10 mV (calibration line) | 3038.5931 | 3.41 | 3.41 | Identical | On | On | Same | ¦12.030V 3.410A0¦ | ¦12.030V 3.410A ¦
623 | 3043.4783 | 12283.4783 | 12.27 | 12.26 | Within 10 mV (calibration line) | 3043.4783 | 3.42 | 3.42 | Identical | On | On | Same | ¦12.030V 3.420A0¦ | ¦12.030V 3.420A ¦
624 | 3048.3635 | 12303.1949 | 12.29 | 12.28 | Within 10 mV (calibration line) | 3048.3635 | 3.42 | 3.42 | Identical | On | On | Same | ¦12.030V 3.420A0¦ | ¦12.030V 3.420A ¦
625 | 3053.2487 | 12322.9116 | 12.31 | 12.3 | Within 10 mV (calibration line) | 3053.2487 | 3.43 | 3.43 | Identical | On | On | Same | ¦12.030V 3.430A0¦ | ¦12.030V 3.430A ¦
626 | 3058.1339 | 12342.6282 | 12.33 | 12.32 | Within 10 mV (calibration line) | 3058.1339 | 3.43 | 3.43 | Identical | On | On | Same | ¦12.030V 3.430A0¦ | ¦12.030V 3.430A ¦
627 | 3063.0191 | 12362.3449 | 12.35 | 12.34 | Within 10 mV (calibration line) | 3063.0191 | 3.44 | 3.44 | Identical | On | On | Same | ¦12.030V 3.440A0¦ | ¦12.030V 3.440A ¦
628 | 3067.9043 | 12382.0616 | 12.37 | 12.36 | Within 10 mV (calibration line) | 3067.9043 | 3.44 | 3.45 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 3.440A0¦ | ¦12.030V 3.450A ¦
629 | 3072.7894 | 12401.7782 | 12.39 | 12.38 | Within 10 mV (calibration line) | 3072.7894 | 3.45 | 3.45 | Identical | On | On | Same | ¦12.030V 3.450A0¦ | ¦12.030V 3.450A ¦
630 | 3077.6746 | 12421.4949 | 12.41 | 12.4 | Within 10 mV (calibration line) | 3077.6746 | 3.46 | 3.46 | Identical | On | On | Same | ¦12.030V 3.460A0¦ | ¦12.030V 3.460A ¦
631 | 3082.5598 | 12441.2115 | 12.43 | 12.42 | Within 10 mV (calibration line) | 3082.5598 | 3.46 | 3.46 | Identical | On | On | Same | ¦12.030V 3.460A0¦ | ¦12.030V 3.460A ¦
632 | 3087.445 | 12460.9282 | 12.45 | 12.44 | Within 10 mV (calibration line) | 3087.445 | 3.47 | 3.47 | Identical | On | On | Same | ¦12.030V 3.470A0¦ | ¦12.030V 3.470A ¦
633 | 3092.3302 | 12480.6448 | 12.47 | 12.46 | Within 10 mV (calibration line) | 3092.3302 | 3.47 | 3.47 | Identical | On | On | Same | ¦12.030V 3.470A0¦ | ¦12.030V 3.470A ¦
634 | 3097.2154 | 12500.3615 | 12.49 | 12.48 | Within 10 mV (calibration line) | 3097.2154 | 3.48 | 3.48 | Identical | On | On | Same | ¦12.030V 3.480A0¦ | ¦12.030V 3.480A ¦
635 | 3102.1006 | 12520.0782 | 12.51 | 12.5 | Within 10 mV (calibration line) | 3102.1006 | 3.48 | 3.48 | Identical | On | On | Same | ¦12.030V 3.480A0¦ | ¦12.030V 3.480A ¦
636 | 3106.9858 | 12539.7948 | 12.52 | 12.52 | Identical | 3106.9858 | 3.49 | 3.49 | Identical | On | On | Same | ¦12.030V 3.490A0¦ | ¦12.030V 3.490A ¦
637 | 3111.871 | 12559.5115 | 12.54 | 12.54 | Identical | 3111.871 | 3.49 | 3.5 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 3.490A0¦ | ¦12.030V 3.500A ¦
638 | 3116.7562 | 12579.2281 | 12.56 | 12.56 | Identical | 3116.7562 | 3.5 | 3.5 | Identical | On | On | Same | ¦12.030V 3.500A0¦ | ¦12.030V 3.500A ¦
639 | 3121.6414 | 12598.9448 | 12.58 | 12.58 | Identical | 3121.6414 | 3.51 | 3.51 | Identical | On | On | Same | ¦12.030V 3.510A0¦ | ¦12.030V 3.510A ¦
640 | 3126.5266 | 12618.6615 | 12.6 | 12.6 | Identical | 3126.5266 | 3.51 | 3.51 | Identical | On | On | Same | ¦12.030V 3.510A0¦ | ¦12.030V 3.510A ¦
641 | 3131.4118 | 12638.3781 | 12.62 | 12.62 | Identical | 3131.4118 | 3.52 | 3.52 | Identical | On | On | Same | ¦12.030V 3.520A0¦ | ¦12.030V 3.520A ¦
642 | 3136.297 | 12658.0948 | 12.64 | 12.64 | Identical | 3136.297 | 3.52 | 3.52 | Identical | On | On | Same | ¦12.030V 3.520A0¦ | ¦12.030V 3.520A ¦
643 | 3141.1822 | 12677.8114 | 12.66 | 12.66 | Identical | 3141.1822 | 3.53 | 3.53 | Identical | On | On | Same | ¦12.030V 3.530A0¦ | ¦12.030V 3.530A ¦
644 | 3146.0674 | 12697.5281 | 12.68 | 12.67 | Within 10 mV (calibration line) | 3146.0674 | 3.53 | 3.53 | Identical | On | On | Same | ¦12.030V 3.530A0¦ | ¦12.030V 3.530A ¦
645 | 3150.9526 | 12717.2447 | 12.7 | 12.69 | Within 10 mV (calibration line) | 3150.9526 | 3.54 | 3.54 | Identical | On | On | Same | ¦12.030V 3.540A0¦ | ¦12.030V 3.540A ¦
646 | 3155.8378 | 12736.9614 | 12.72 | 12.71 | Within 10 mV (calibration line) | 3155.8378 | 3.54 | 3.54 | Identical | On | On | Same | ¦12.030V 3.540A0¦ | ¦12.030V 3.540A ¦
647 | 3160.723 | 12756.6781 | 12.74 | 12.73 | Within 10 mV (calibration line) | 3160.723 | 3.55 | 3.55 | Identical | On | On | Same | ¦12.030V 3.550A0¦ | ¦12.030V 3.550A ¦
648 | 3165.6082 | 12776.3947 | 12.76 | 12.75 | Within 10 mV (calibration line) | 3165.6082 | 3.55 | 3.56 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 3.550A0¦ | ¦12.030V 3.560A ¦
649 | 3170.4934 | 12796.1114 | 12.78 | 12.77 | Within 10 mV (calibration line) | 3170.4934 | 3.56 | 3.56 | Identical | On | On | Same | ¦12.030V 3.560A0¦ | ¦12.030V 3.560A ¦
650 | 3175.3786 | 12815.828 | 12.8 | 12.79 | Within 10 mV (calibration line) | 3175.3786 | 3.57 | 3.57 | Identical | On | On | Same | ¦12.030V 3.570A0¦ | ¦12.030V 3.570A ¦
651 | 3180.2638 | 12835.5447 | 12.82 | 12.81 | Within 10 mV (calibration line) | 3180.2638 | 3.57 | 3.57 | Identical | On | On | Same | ¦12.030V 3.570A0¦ | ¦12.030V 3.570A ¦
652 | 3185.149 | 12855.2614 | 12.84 | 12.83 | Within 10 mV (calibration line) | 3185.149 | 3.58 | 3.58 | Identical | On | On | Same | ¦12.030V 3.580A0¦ | ¦12.030V 3.580A ¦
653 | 3190.0342 | 12874.978 | 12.86 | 12.85 | Within 10 mV (calibration line) | 3190.0342 | 3.58 | 3.58 | Identical | On | On | Same | ¦12.030V 3.580A0¦ | ¦12.030V 3.580A ¦
654 | 3194.9194 | 12894.6947 | 12.88 | 12.87 | Within 10 mV (calibration line) | 3194.9194 | 3.59 | 3.59 | Identical | On | On | Same | ¦12.030V 3.590A0¦ | ¦12.030V 3.590A ¦
655 | 3199.8046 | 12914.4113 | 12.9 | 12.89 | Within 10 mV (calibration line) | 3199.8046 | 3.59 | 3.59 | Identical | On | On | Same | ¦12.030V 3.590A0¦ | ¦12.030V 3.590A ¦
656 | 3204.6898 | 12934.128 | 12.92 | 12.91 | Within 10 mV (calibration line) | 3204.6898 | 3.6 | 3.6 | Identical | On | On | Same | ¦12.030V 3.600A0¦ | ¦12.030V 3.600A ¦
657 | 3209.575 | 12953.8447 | 12.94 | 12.93 | Within 10 mV (calibration line) | 3209.575 | 3.6 | 3.6 | Identical | On | On | Same | ¦12.030V 3.600A0¦ | ¦12.030V 3.600A ¦
658 | 3214.4602 | 12973.5613 | 12.96 | 12.95 | Within 10 mV (calibration line) | 3214.4602 | 3.61 | 3.61 | Identical | On | On | Same | ¦12.030V 3.610A0¦ | ¦12.030V 3.610A ¦
659 | 3219.3454 | 12993.278 | 12.98 | 12.97 | Within 10 mV (calibration line) | 3219.3454 | 3.61 | 3.61 | Identical | On | On | Same | ¦12.030V 3.610A0¦ | ¦12.030V 3.610A ¦
660 | 3224.2306 | 13012.9946 | 13 | 12.99 | Within 10 mV (calibration line) | 3224.2306 | 3.62 | 3.62 | Identical | On | On | Same | ¦12.030V 3.620A0¦ | ¦12.030V 3.620A ¦
661 | 3229.1158 | 13032.7113 | 13.01 | 13.01 | Identical | 3229.1158 | 3.63 | 3.63 | Identical | On | On | Same | ¦12.030V 3.630A0¦ | ¦12.030V 3.630A ¦
662 | 3234.001 | 13052.4279 | 13.03 | 13.03 | Identical | 3234.001 | 3.63 | 3.63 | Identical | On | On | Same | ¦12.030V 3.630A0¦ | ¦12.030V 3.630A ¦
663 | 3238.8862 | 13072.1446 | 13.05 | 13.05 | Identical | 3238.8862 | 3.64 | 3.64 | Identical | On | On | Same | ¦12.030V 3.640A0¦ | ¦12.030V 3.640A ¦
664 | 3243.7714 | 13091.8613 | 13.07 | 13.07 | Identical | 3243.7714 | 3.64 | 3.64 | Identical | On | On | Same | ¦12.030V 3.640A0¦ | ¦12.030V 3.640A ¦
665 | 3248.6566 | 13111.5779 | 13.09 | 13.09 | Identical | 3248.6566 | 3.65 | 3.65 | Identical | On | On | Same | ¦12.030V 3.650A0¦ | ¦12.030V 3.650A ¦
666 | 3253.5418 | 13131.2946 | 13.11 | 13.11 | Identical | 3253.5418 | 3.65 | 3.65 | Identical | On | On | Same | ¦12.030V 3.650A0¦ | ¦12.030V 3.650A ¦
667 | 3258.427 | 13151.0112 | 13.13 | 13.13 | Identical | 3258.427 | 3.66 | 3.66 | Identical | On | On | Same | ¦12.030V 3.660A0¦ | ¦12.030V 3.660A ¦
668 | 3263.3122 | 13170.7279 | 13.15 | 13.15 | Identical | 3263.3122 | 3.66 | 3.66 | Identical | On | On | Same | ¦12.030V 3.660A0¦ | ¦12.030V 3.660A ¦
669 | 3268.1974 | 13190.4446 | 13.17 | 13.17 | Identical | 3268.1974 | 3.67 | 3.67 | Identical | On | On | Same | ¦12.030V 3.670A0¦ | ¦12.030V 3.670A ¦
670 | 3273.0826 | 13210.1612 | 13.19 | 13.19 | Identical | 3273.0826 | 3.68 | 3.68 | Identical | On | On | Same | ¦12.030V 3.680A0¦ | ¦12.030V 3.680A ¦
671 | 3277.9678 | 13229.8779 | 13.2 | 13.21 | Within 10 mV (calibration line) | 3277.9678 | 3.68 | 3.68 | Identical | On | On | Same | ¦12.030V 3.680A0¦ | ¦12.030V 3.680A ¦
672 | 3282.853 | 13249.5945 | 13.22 | 13.22 | Identical | 3282.853 | 3.69 | 3.69 | Identical | On | On | Same | ¦12.030V 3.690A0¦ | ¦12.030V 3.690A ¦
673 | 3287.7382 | 13269.3112 | 13.24 | 13.24 | Identical | 3287.7382 | 3.69 | 3.69 | Identical | On | On | Same | ¦12.030V 3.690A0¦ | ¦12.030V 3.690A ¦
674 | 3292.6234 | 13289.0278 | 13.26 | 13.26 | Identical | 3292.6234 | 3.7 | 3.7 | Identical | On | On | Same | ¦12.030V 3.700A0¦ | ¦12.030V 3.700A ¦
675 | 3297.5085 | 13308.7445 | 13.28 | 13.28 | Identical | 3297.5085 | 3.7 | 3.7 | Identical | On | On | Same | ¦12.030V 3.700A0¦ | ¦12.030V 3.700A ¦
676 | 3302.3937 | 13328.4612 | 13.3 | 13.3 | Identical | 3302.3937 | 3.71 | 3.71 | Identical | On | On | Same | ¦12.030V 3.710A0¦ | ¦12.030V 3.710A ¦
677 | 3307.2789 | 13348.1778 | 13.32 | 13.32 | Identical | 3307.2789 | 3.71 | 3.71 | Identical | On | On | Same | ¦12.030V 3.710A0¦ | ¦12.030V 3.710A ¦
678 | 3312.1641 | 13367.8945 | 13.34 | 13.34 | Identical | 3312.1641 | 3.72 | 3.72 | Identical | On | On | Same | ¦12.030V 3.720A0¦ | ¦12.030V 3.720A ¦
679 | 3317.0493 | 13387.6111 | 13.36 | 13.36 | Identical | 3317.0493 | 3.72 | 3.72 | Identical | On | On | Same | ¦12.030V 3.720A0¦ | ¦12.030V 3.720A ¦
680 | 3321.9345 | 13407.3278 | 13.38 | 13.38 | Identical | 3321.9345 | 3.73 | 3.73 | Identical | On | On | Same | ¦12.030V 3.730A0¦ | ¦12.030V 3.730A ¦
681 | 3326.8197 | 13427.0445 | 13.4 | 13.4 | Identical | 3326.8197 | 3.73 | 3.74 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 3.730A0¦ | ¦12.030V 3.740A ¦
682 | 3331.7049 | 13446.7611 | 13.42 | 13.42 | Identical | 3331.7049 | 3.74 | 3.74 | Identical | On | On | Same | ¦12.030V 3.740A0¦ | ¦12.030V 3.740A ¦
683 | 3336.5901 | 13466.4778 | 13.44 | 13.44 | Identical | 3336.5901 | 3.75 | 3.75 | Identical | On | On | Same | ¦12.030V 3.750A0¦ | ¦12.030V 3.750A ¦
684 | 3341.4753 | 13486.1944 | 13.46 | 13.46 | Identical | 3341.4753 | 3.75 | 3.75 | Identical | On | On | Same | ¦12.030V 3.750A0¦ | ¦12.030V 3.750A ¦
685 | 3346.3605 | 13505.9111 | 13.48 | 13.48 | Identical | 3346.3605 | 3.76 | 3.76 | Identical | On | On | Same | ¦12.030V 3.760A0¦ | ¦12.030V 3.760A ¦
686 | 3351.2457 | 13525.6277 | 13.5 | 13.5 | Identical | 3351.2457 | 3.76 | 3.76 | Identical | On | On | Same | ¦12.030V 3.760A0¦ | ¦12.030V 3.760A ¦
687 | 3356.1309 | 13545.3444 | 13.52 | 13.52 | Identical | 3356.1309 | 3.77 | 3.77 | Identical | On | On | Same | ¦12.030V 3.770A0¦ | ¦12.030V 3.770A ¦
688 | 3361.0161 | 13565.0611 | 13.54 | 13.54 | Identical | 3361.0161 | 3.77 | 3.77 | Identical | On | On | Same | ¦12.030V 3.770A0¦ | ¦12.030V 3.770A ¦
689 | 3365.9013 | 13584.7777 | 13.56 | 13.56 | Identical | 3365.9013 | 3.78 | 3.78 | Identical | On | On | Same | ¦12.030V 3.780A0¦ | ¦12.030V 3.780A ¦
690 | 3370.7865 | 13604.4944 | 13.58 | 13.58 | Identical | 3370.7865 | 3.78 | 3.78 | Identical | On | On | Same | ¦12.030V 3.780A0¦ | ¦12.030V 3.780A ¦
691 | 3375.6717 | 13624.211 | 13.6 | 13.6 | Identical | 3375.6717 | 3.79 | 3.79 | Identical | On | On | Same | ¦12.030V 3.790A0¦ | ¦12.030V 3.790A ¦
692 | 3380.5569 | 13643.9277 | 13.62 | 13.62 | Identical | 3380.5569 | 3.79 | 3.8 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 3.790A0¦ | ¦12.030V 3.800A ¦
693 | 3385.4421 | 13663.6444 | 13.64 | 13.64 | Identical | 3385.4421 | 3.8 | 3.8 | Identical | On | On | Same | ¦12.030V 3.800A0¦ | ¦12.030V 3.800A ¦
694 | 3390.3273 | 13683.361 | 13.66 | 13.66 | Identical | 3390.3273 | 3.81 | 3.81 | Identical | On | On | Same | ¦12.030V 3.810A0¦ | ¦12.030V 3.810A ¦
695 | 3395.2125 | 13703.0777 | 13.68 | 13.68 | Identical | 3395.2125 | 3.81 | 3.81 | Identical | On | On | Same | ¦12.030V 3.810A0¦ | ¦12.030V 3.810A ¦
696 | 3400.0977 | 13722.7943 | 13.7 | 13.7 | Identical | 3400.0977 | 3.82 | 3.82 | Identical | On | On | Same | ¦12.030V 3.820A0¦ | ¦12.030V 3.820A ¦
697 | 3404.9829 | 13742.511 | 13.72 | 13.72 | Identical | 3404.9829 | 3.82 | 3.82 | Identical | On | On | Same | ¦12.030V 3.820A0¦ | ¦12.030V 3.820A ¦
698 | 3409.8681 | 13762.2277 | 13.74 | 13.74 | Identical | 3409.8681 | 3.83 | 3.83 | Identical | On | On | Same | ¦12.030V 3.830A0¦ | ¦12.030V 3.830A ¦
699 | 3414.7533 | 13781.9443 | 13.76 | 13.75 | Within 10 mV (calibration line) | 3414.7533 | 3.83 | 3.83 | Identical | On | On | Same | ¦12.030V 3.830A0¦ | ¦12.030V 3.830A ¦
700 | 3419.6385 | 13801.661 | 13.78 | 13.77 | Within 10 mV (calibration line) | 3419.6385 | 3.84 | 3.84 | Identical | On | On | Same | ¦12.030V 3.840A0¦ | ¦12.030V 3.840A ¦
701 | 3424.5237 | 13821.3776 | 13.8 | 13.79 | Within 10 mV (calibration line) | 3424.5237 | 3.84 | 3.84 | Identical | On | On | Same | ¦12.030V 3.840A0¦ | ¦12.030V 3.840A ¦
702 | 3429.4089 | 13841.0943 | 13.82 | 13.81 | Within 10 mV (calibration line) | 3429.4089 | 3.85 | 3.85 | Identical | On | On | Same | ¦12.030V 3.850A0¦ | ¦12.030V 3.850A ¦
703 | 3434.2941 | 13860.8109 | 13.84 | 13.83 | Within 10 mV (calibration line) | 3434.2941 | 3.86 | 3.86 | Identical | On | On | Same | ¦12.030V 3.860A0¦ | ¦12.030V 3.860A ¦
704 | 3439.1793 | 13880.5276 | 13.86 | 13.85 | Within 10 mV (calibration line) | 3439.1793 | 3.86 | 3.86 | Identical | On | On | Same | ¦12.030V 3.860A0¦ | ¦12.030V 3.860A ¦
705 | 3444.0645 | 13900.2443 | 13.88 | 13.87 | Within 10 mV (calibration line) | 3444.0645 | 3.87 | 3.87 | Identical | On | On | Same | ¦12.030V 3.870A0¦ | ¦12.030V 3.870A ¦
706 | 3448.9497 | 13919.9609 | 13.89 | 13.89 | Identical | 3448.9497 | 3.87 | 3.87 | Identical | On | On | Same | ¦12.030V 3.870A0¦ | ¦12.030V 3.870A ¦
707 | 3453.8349 | 13939.6776 | 13.91 | 13.91 | Identical | 3453.8349 | 3.88 | 3.88 | Identical | On | On | Same | ¦12.030V 3.880A0¦ | ¦12.030V 3.880A ¦
708 | 3458.7201 | 13959.3942 | 13.93 | 13.93 | Identical | 3458.7201 | 3.88 | 3.88 | Identical | On | On | Same | ¦12.030V 3.880A0¦ | ¦12.030V 3.880A ¦
709 | 3463.6053 | 13979.1109 | 13.95 | 13.95 | Identical | 3463.6053 | 3.89 | 3.89 | Identical | On | On | Same | ¦12.030V 3.890A0¦ | ¦12.030V 3.890A ¦
710 | 3468.4905 | 13998.8276 | 13.97 | 13.97 | Identical | 3468.4905 | 3.89 | 3.89 | Identical | On | On | Same | ¦12.030V 3.890A0¦ | ¦12.030V 3.890A ¦
711 | 3473.3757 | 14018.5442 | 13.99 | 13.99 | Identical | 3473.3757 | 3.9 | 3.9 | Identical | On | On | Same | ¦12.030V 3.900A0¦ | ¦12.030V 3.900A ¦
712 | 3478.2609 | 14038.2609 | 14.01 | 14.01 | Identical | 3478.2609 | 3.9 | 3.9 | Identical | On | On | Same | ¦12.030V 3.900A0¦ | ¦12.030V 3.900A ¦
713 | 3483.1461 | 14057.9775 | 14.03 | 14.03 | Identical | 3483.1461 | 3.91 | 3.91 | Identical | On | On | Same | ¦12.030V 3.910A0¦ | ¦12.030V 3.910A ¦
714 | 3488.0313 | 14077.6942 | 14.05 | 14.05 | Identical | 3488.0313 | 3.92 | 3.92 | Identical | On | On | Same | ¦12.030V 3.920A0¦ | ¦12.030V 3.920A ¦
715 | 3492.9165 | 14097.4108 | 14.07 | 14.07 | Identical | 3492.9165 | 3.92 | 3.92 | Identical | On | On | Same | ¦12.030V 3.920A0¦ | ¦12.030V 3.920A ¦
716 | 3497.8017 | 14117.1275 | 14.09 | 14.09 | Identical | 3497.8017 | 3.93 | 3.93 | Identical | On | On | Same | ¦12.030V 3.930A0¦ | ¦12.030V 3.930A ¦
717 | 3502.6869 | 14136.8442 | 14.11 | 14.11 | Identical | 3502.6869 | 3.93 | 3.93 | Identical | On | On | Same | ¦12.030V 3.930A0¦ | ¦12.030V 3.930A ¦
718 | 3507.5721 | 14156.5608 | 14.13 | 14.13 | Identical | 3507.5721 | 3.94 | 3.94 | Identical | On | On | Same | ¦12.030V 3.940A0¦ | ¦12.030V 3.940A ¦
719 | 3512.4573 | 14176.2775 | 14.15 | 14.15 | Identical | 3512.4573 | 3.94 | 3.94 | Identical | On | On | Same | ¦12.030V 3.940A0¦ | ¦12.030V 3.940A ¦
720 | 3517.3425 | 14195.9941 | 14.17 | 14.17 | Identical | 3517.3425 | 3.95 | 3.95 | Identical | On | On | Same | ¦12.030V 3.950A0¦ | ¦12.030V 3.950A ¦
721 | 3522.2277 | 14215.7108 | 14.19 | 14.19 | Identical | 3522.2277 | 3.95 | 3.95 | Identical | On | On | Same | ¦12.030V 3.950A0¦ | ¦12.030V 3.950A ¦
722 | 3527.1128 | 14235.4275 | 14.21 | 14.21 | Identical | 3527.1128 | 3.96 | 3.96 | Identical | On | On | Same | ¦12.030V 3.960A0¦ | ¦12.030V 3.960A ¦
723 | 3531.998 | 14255.1441 | 14.23 | 14.23 | Identical | 3531.998 | 3.96 | 3.96 | Identical | On | On | Same | ¦12.030V 3.960A0¦ | ¦12.030V 3.960A ¦
724 | 3536.8832 | 14274.8608 | 14.25 | 14.25 | Identical | 3536.8832 | 3.97 | 3.97 | Identical | On | On | Same | ¦12.030V 3.970A0¦ | ¦12.030V 3.970A ¦
725 | 3541.7684 | 14294.5774 | 14.27 | 14.27 | Identical | 3541.7684 | 3.97 | 3.98 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 3.970A0¦ | ¦12.030V 3.980A ¦
726 | 3546.6536 | 14314.2941 | 14.29 | 14.28 | Within 10 mV (calibration line) | 3546.6536 | 3.98 | 3.98 | Identical | On | On | Same | ¦12.030V 3.980A0¦ | ¦12.030V 3.980A ¦
727 | 3551.5388 | 14334.0107 | 14.31 | 14.3 | Within 10 mV (calibration line) | 3551.5388 | 3.99 | 3.99 | Identical | On | On | Same | ¦12.030V 3.990A0¦ | ¦12.030V 3.990A ¦
728 | 3556.424 | 14353.7274 | 14.33 | 14.32 | Within 10 mV (calibration line) | 3556.424 | 3.99 | 3.99 | Identical | On | On | Same | ¦12.030V 3.990A0¦ | ¦12.030V 3.990A ¦
729 | 3561.3092 | 14373.4441 | 14.35 | 14.34 | Within 10 mV (calibration line) | 3561.3092 | 4 | 4 | Identical | On | On | Same | ¦12.030V 4.000A0¦ | ¦12.030V 4.000A ¦
730 | 3566.1944 | 14393.1607 | 14.37 | 14.36 | Within 10 mV (calibration line) | 3566.1944 | 4 | 4 | Identical | On | On | Same | ¦12.030V 4.000A0¦ | ¦12.030V 4.000A ¦
731 | 3571.0796 | 14412.8774 | 14.39 | 14.38 | Within 10 mV (calibration line) | 3571.0796 | 4.01 | 4.01 | Identical | On | On | Same | ¦12.030V 4.010A0¦ | ¦12.030V 4.010A ¦
732 | 3575.9648 | 14432.594 | 14.41 | 14.4 | Within 10 mV (calibration line) | 3575.9648 | 4.01 | 4.02 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 4.010A0¦ | ¦12.030V 4.020A ¦
733 | 3580.85 | 14452.3107 | 14.43 | 14.42 | Within 10 mV (calibration line) | 3580.85 | 4.02 | 4.02 | Identical | On | On | Same | ¦12.030V 4.020A0¦ | ¦12.030V 4.020A ¦
734 | 3585.7352 | 14472.0274 | 14.45 | 14.44 | Within 10 mV (calibration line) | 3585.7352 | 4.03 | 4.03 | Identical | On | On | Same | ¦12.030V 4.030A0¦ | ¦12.030V 4.030A ¦
735 | 3590.6204 | 14491.744 | 14.47 | 14.46 | Within 10 mV (calibration line) | 3590.6204 | 4.03 | 4.03 | Identical | On | On | Same | ¦12.030V 4.030A0¦ | ¦12.030V 4.030A ¦
736 | 3595.5056 | 14511.4607 | 14.49 | 14.48 | Within 10 mV (calibration line) | 3595.5056 | 4.04 | 4.04 | Identical | On | On | Same | ¦12.030V 4.040A0¦ | ¦12.030V 4.040A ¦
737 | 3600.3908 | 14531.1773 | 14.51 | 14.5 | Within 10 mV (calibration line) | 3600.3908 | 4.04 | 4.04 | Identical | On | On | Same | ¦12.030V 4.040A0¦ | ¦12.030V 4.040A ¦
738 | 3605.276 | 14550.894 | 14.53 | 14.52 | Within 10 mV (calibration line) | 3605.276 | 4.05 | 4.05 | Identical | On | On | Same | ¦12.030V 4.050A0¦ | ¦12.030V 4.050A ¦
739 | 3610.1612 | 14570.6106 | 14.55 | 14.54 | Within 10 mV (calibration line) | 3610.1612 | 4.05 | 4.05 | Identical | On | On | Same | ¦12.030V 4.050A0¦ | ¦12.030V 4.050A ¦
740 | 3615.0464 | 14590.3273 | 14.57 | 14.56 | Within 10 mV (calibration line) | 3615.0464 | 4.06 | 4.06 | Identical | On | On | Same | ¦12.030V 4.060A0¦ | ¦12.030V 4.060A ¦
741 | 3619.9316 | 14610.044 | 14.59 | 14.58 | Within 10 mV (calibration line) | 3619.9316 | 4.06 | 4.07 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 4.060A0¦ | ¦12.030V 4.070A ¦
742 | 3624.8168 | 14629.7606 | 14.6 | 14.6 | Identical | 3624.8168 | 4.07 | 4.07 | Identical | On | On | Same | ¦12.030V 4.070A0¦ | ¦12.030V 4.070A ¦
743 | 3629.702 | 14649.4773 | 14.62 | 14.62 | Identical | 3629.702 | 4.08 | 4.08 | Identical | On | On | Same | ¦12.030V 4.080A0¦ | ¦12.030V 4.080A ¦
744 | 3634.5872 | 14669.1939 | 14.64 | 14.64 | Identical | 3634.5872 | 4.08 | 4.08 | Identical | On | On | Same | ¦12.030V 4.080A0¦ | ¦12.030V 4.080A ¦
745 | 3639.4724 | 14688.9106 | 14.66 | 14.66 | Identical | 3639.4724 | 4.09 | 4.09 | Identical | On | On | Same | ¦12.030V 4.090A0¦ | ¦12.030V 4.090A ¦
746 | 3644.3576 | 14708.6273 | 14.68 | 14.68 | Identical | 3644.3576 | 4.09 | 4.09 | Identical | On | On | Same | ¦12.030V 4.090A0¦ | ¦12.030V 4.090A ¦
747 | 3649.2428 | 14728.3439 | 14.7 | 14.7 | Identical | 3649.2428 | 4.1 | 4.1 | Identical | On | On | Same | ¦12.030V 4.100A0¦ | ¦12.030V 4.100A ¦
748 | 3654.128 | 14748.0606 | 14.72 | 14.72 | Identical | 3654.128 | 4.1 | 4.1 | Identical | On | On | Same | ¦12.030V 4.100A0¦ | ¦12.030V 4.100A ¦
749 | 3659.0132 | 14767.7772 | 14.74 | 14.74 | Identical | 3659.0132 | 4.11 | 4.11 | Identical | On | On | Same | ¦12.030V 4.110A0¦ | ¦12.030V 4.110A ¦
750 | 3663.8984 | 14787.4939 | 14.76 | 14.76 | Identical | 3663.8984 | 4.11 | 4.11 | Identical | On | On | Same | ¦12.030V 4.110A0¦ | ¦12.030V 4.110A ¦
751 | 3668.7836 | 14807.2106 | 14.78 | 14.78 | Identical | 3668.7836 | 4.12 | 4.12 | Identical | On | On | Same | ¦12.030V 4.120A0¦ | ¦12.030V 4.120A ¦
752 | 3673.6688 | 14826.9272 | 14.8 | 14.8 | Identical | 3673.6688 | 4.12 | 4.13 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 4.120A0¦ | ¦12.030V 4.130A ¦
753 | 3678.554 | 14846.6439 | 14.82 | 14.81 | Within 10 mV (calibration line) | 3678.554 | 4.13 | 4.13 | Identical | On | On | Same | ¦12.030V 4.130A0¦ | ¦12.030V 4.130A ¦
754 | 3683.4392 | 14866.3605 | 14.84 | 14.83 | Within 10 mV (calibration line) | 3683.4392 | 4.14 | 4.14 | Identical | On | On | Same | ¦12.030V 4.140A0¦ | ¦12.030V 4.140A ¦
755 | 3688.3244 | 14886.0772 | 14.86 | 14.85 | Within 10 mV (calibration line) | 3688.3244 | 4.14 | 4.14 | Identical | On | On | Same | ¦12.030V 4.140A0¦ | ¦12.030V 4.140A ¦
756 | 3693.2096 | 14905.7938 | 14.88 | 14.87 | Within 10 mV (calibration line) | 3693.2096 | 4.15 | 4.15 | Identical | On | On | Same | ¦12.030V 4.150A0¦ | ¦12.030V 4.150A ¦
757 | 3698.0948 | 14925.5105 | 14.9 | 14.89 | Within 10 mV (calibration line) | 3698.0948 | 4.15 | 4.15 | Identical | On | On | Same | ¦12.030V 4.150A0¦ | ¦12.030V 4.150A ¦
758 | 3702.98 | 14945.2272 | 14.92 | 14.91 | Within 10 mV (calibration line) | 3702.98 | 4.16 | 4.16 | Identical | On | On | Same | ¦12.030V 4.160A0¦ | ¦12.030V 4.160A ¦
759 | 3707.8652 | 14964.9438 | 14.94 | 14.93 | Within 10 mV (calibration line) | 3707.8652 | 4.16 | 4.16 | Identical | On | On | Same | ¦12.030V 4.160A0¦ | ¦12.030V 4.160A ¦
760 | 3712.7504 | 14984.6605 | 14.96 | 14.95 | Within 10 mV (calibration line) | 3712.7504 | 4.17 | 4.17 | Identical | On | On | Same | ¦12.030V 4.170A0¦ | ¦12.030V 4.170A ¦
761 | 3717.6356 | 15004.3771 | 14.98 | 14.97 | Within 10 mV (calibration line) | 3717.6356 | 4.17 | 4.17 | Identical | On | On | Same | ¦12.030V 4.170A0¦ | ¦12.030V 4.170A ¦
762 | 3722.5208 | 15024.0938 | 15 | 14.99 | Within 10 mV (calibration line) | 3722.5208 | 4.18 | 4.18 | Identical | On | On | Same | ¦12.030V 4.180A0¦ | ¦12.030V 4.180A ¦
763 | 3727.406 | 15043.8105 | 15.01 | 15.01 | Identical | 3727.406 | 4.18 | 4.18 | Identical | On | On | Same | ¦12.030V 4.180A0¦ | ¦12.030V 4.180A ¦
764 | 3732.2912 | 15063.5271 | 15.03 | 15.03 | Identical | 3732.2912 | 4.19 | 4.19 | Identical | On | On | Same | ¦12.030V 4.190A0¦ | ¦12.030V 4.190A ¦
765 | 3737.1764 | 15083.2438 | 15.05 | 15.05 | Identical | 3737.1764 | 4.2 | 4.2 | Identical | On | On | Same | ¦12.030V 4.200A0¦ | ¦12.030V 4.200A ¦
766 | 3742.0616 | 15102.9604 | 15.07 | 15.07 | Identical | 3742.0616 | 4.2 | 4.2 | Identical | On | On | Same | ¦12.030V 4.200A0¦ | ¦12.030V 4.200A ¦
767 | 3746.9468 | 15122.6771 | 15.09 | 15.09 | Identical | 3746.9468 | 4.21 | 4.21 | Identical | On | On | Same | ¦12.030V 4.210A0¦ | ¦12.030V 4.210A ¦
768 | 3751.8319 | 15142.3937 | 15.11 | 15.11 | Identical | 3751.8319 | 4.21 | 4.21 | Identical | On | On | Same | ¦12.030V 4.210A0¦ | ¦12.030V 4.210A ¦
769 | 3756.7171 | 15162.1104 | 15.13 | 15.13 | Identical | 3756.7171 | 4.22 | 4.22 | Identical | On | On | Same | ¦12.030V 4.220A0¦ | ¦12.030V 4.220A ¦
770 | 3761.6023 | 15181.8271 | 15.15 | 15.15 | Identical | 3761.6023 | 4.22 | 4.22 | Identical | On | On | Same | ¦12.030V 4.220A0¦ | ¦12.030V 4.220A ¦
771 | 3766.4875 | 15201.5437 | 15.17 | 15.17 | Identical | 3766.4875 | 4.23 | 4.23 | Identical | On | On | Same | ¦12.030V 4.230A0¦ | ¦12.030V 4.230A ¦
772 | 3771.3727 | 15221.2604 | 15.19 | 15.19 | Identical | 3771.3727 | 4.23 | 4.23 | Identical | On | On | Same | ¦12.030V 4.230A0¦ | ¦12.030V 4.230A ¦
773 | 3776.2579 | 15240.977 | 15.21 | 15.21 | Identical | 3776.2579 | 4.24 | 4.24 | Identical | On | On | Same | ¦12.030V 4.240A0¦ | ¦12.030V 4.240A ¦
774 | 3781.1431 | 15260.6937 | 15.23 | 15.23 | Identical | 3781.1431 | 4.25 | 4.25 | Identical | On | On | Same | ¦12.030V 4.250A0¦ | ¦12.030V 4.250A ¦
775 | 3786.0283 | 15280.4104 | 15.25 | 15.25 | Identical | 3786.0283 | 4.25 | 4.25 | Identical | On | On | Same | ¦12.030V 4.250A0¦ | ¦12.030V 4.250A ¦
776 | 3790.9135 | 15300.127 | 15.27 | 15.27 | Identical | 3790.9135 | 4.26 | 4.26 | Identical | On | On | Same | ¦12.030V 4.260A0¦ | ¦12.030V 4.260A ¦
777 | 3795.7987 | 15319.8437 | 15.28 | 15.29 | Within 10 mV (calibration line) | 3795.7987 | 4.26 | 4.26 | Identical | On | On | Same | ¦12.030V 4.260A0¦ | ¦12.030V 4.260A ¦
778 | 3800.6839 | 15339.5603 | 15.3 | 15.31 | Within 10 mV (calibration line) | 3800.6839 | 4.27 | 4.27 | Identical | On | On | Same | ¦12.030V 4.270A0¦ | ¦12.030V 4.270A ¦
779 | 3805.5691 | 15359.277 | 15.32 | 15.33 | Within 10 mV (calibration line) | 3805.5691 | 4.27 | 4.27 | Identical | On | On | Same | ¦12.030V 4.270A0¦ | ¦12.030V 4.270A ¦
780 | 3810.4543 | 15378.9936 | 15.34 | 15.35 | Within 10 mV (calibration line) | 3810.4543 | 4.28 | 4.28 | Identical | On | On | Same | ¦12.030V 4.280A0¦ | ¦12.030V 4.280A ¦
781 | 3815.3395 | 15398.7103 | 15.36 | 15.36 | Identical | 3815.3395 | 4.28 | 4.28 | Identical | On | On | Same | ¦12.030V 4.280A0¦ | ¦12.030V 4.280A ¦
782 | 3820.2247 | 15418.427 | 15.38 | 15.38 | Identical | 3820.2247 | 4.29 | 4.29 | Identical | On | On | Same | ¦12.030V 4.290A0¦ | ¦12.030V 4.290A ¦
783 | 3825.1099 | 15438.1436 | 15.4 | 15.4 | Identical | 3825.1099 | 4.29 | 4.29 | Identical | On | On | Same | ¦12.030V 4.290A0¦ | ¦12.030V 4.290A ¦
784 | 3829.9951 | 15457.8603 | 15.42 | 15.42 | Identical | 3829.9951 | 4.3 | 4.3 | Identical | On | On | Same | ¦12.030V 4.300A0¦ | ¦12.030V 4.300A ¦
785 | 3834.8803 | 15477.5769 | 15.44 | 15.44 | Identical | 3834.8803 | 4.3 | 4.31 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 4.300A0¦ | ¦12.030V 4.310A ¦
786 | 3839.7655 | 15497.2936 | 15.46 | 15.46 | Identical | 3839.7655 | 4.31 | 4.31 | Identical | On | On | Same | ¦12.030V 4.310A0¦ | ¦12.030V 4.310A ¦
787 | 3844.6507 | 15517.0103 | 15.48 | 15.48 | Identical | 3844.6507 | 4.32 | 4.32 | Identical | On | On | Same | ¦12.030V 4.320A0¦ | ¦12.030V 4.320A ¦
788 | 3849.5359 | 15536.7269 | 15.5 | 15.5 | Identical | 3849.5359 | 4.32 | 4.32 | Identical | On | On | Same | ¦12.030V 4.320A0¦ | ¦12.030V 4.320A ¦
789 | 3854.4211 | 15556.4436 | 15.52 | 15.52 | Identical | 3854.4211 | 4.33 | 4.33 | Identical | On | On | Same | ¦12.030V 4.330A0¦ | ¦12.030V 4.330A ¦
790 | 3859.3063 | 15576.1602 | 15.54 | 15.54 | Identical | 3859.3063 | 4.33 | 4.33 | Identical | On | On | Same | ¦12.030V 4.330A0¦ | ¦12.030V 4.330A ¦
791 | 3864.1915 | 15595.8769 | 15.56 | 15.56 | Identical | 3864.1915 | 4.34 | 4.34 | Identical | On | On | Same | ¦12.030V 4.340A0¦ | ¦12.030V 4.340A ¦
792 | 3869.0767 | 15615.5936 | 15.58 | 15.58 | Identical | 3869.0767 | 4.34 | 4.34 | Identical | On | On | Same | ¦12.030V 4.340A0¦ | ¦12.030V 4.340A ¦
793 | 3873.9619 | 15635.3102 | 15.6 | 15.6 | Identical | 3873.9619 | 4.35 | 4.35 | Identical | On | On | Same | ¦12.030V 4.350A0¦ | ¦12.030V 4.350A ¦
794 | 3878.8471 | 15655.0269 | 15.62 | 15.62 | Identical | 3878.8471 | 4.35 | 4.35 | Identical | On | On | Same | ¦12.030V 4.350A0¦ | ¦12.030V 4.350A ¦
795 | 3883.7323 | 15674.7435 | 15.64 | 15.64 | Identical | 3883.7323 | 4.36 | 4.36 | Identical | On | On | Same | ¦12.030V 4.360A0¦ | ¦12.030V 4.360A ¦
796 | 3888.6175 | 15694.4602 | 15.66 | 15.66 | Identical | 3888.6175 | 4.36 | 4.37 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 4.360A0¦ | ¦12.030V 4.370A ¦
797 | 3893.5027 | 15714.1768 | 15.68 | 15.68 | Identical | 3893.5027 | 4.37 | 4.37 | Identical | On | On | Same | ¦12.030V 4.370A0¦ | ¦12.030V 4.370A ¦
798 | 3898.3879 | 15733.8935 | 15.7 | 15.7 | Identical | 3898.3879 | 4.38 | 4.38 | Identical | On | On | Same | ¦12.030V 4.380A0¦ | ¦12.030V 4.380A ¦
799 | 3903.2731 | 15753.6102 | 15.72 | 15.72 | Identical | 3903.2731 | 4.38 | 4.38 | Identical | On | On | Same | ¦12.030V 4.380A0¦ | ¦12.030V 4.380A ¦
800 | 3908.1583 | 15773.3268 | 15.74 | 15.74 | Identical | 3908.1583 | 4.39 | 4.39 | Identical | On | On | Same | ¦12.030V 4.390A0¦ | ¦12.030V 4.390A ¦
801 | 3913.0435 | 15793.0435 | 15.76 | 15.76 | Identical | 3913.0435 | 4.39 | 4.39 | Identical | On | On | Same | ¦12.030V 4.390A0¦ | ¦12.030V 4.390A ¦
802 | 3917.9287 | 15812.7601 | 15.78 | 15.78 | Identical | 3917.9287 | 4.4 | 4.4 | Identical | On | On | Same | ¦12.030V 4.400A0¦ | ¦12.030V 4.400A ¦
803 | 3922.8139 | 15832.4768 | 15.8 | 15.8 | Identical | 3922.8139 | 4.4 | 4.4 | Identical | On | On | Same | ¦12.030V 4.400A0¦ | ¦12.030V 4.400A ¦
804 | 3927.6991 | 15852.1935 | 15.82 | 15.82 | Identical | 3927.6991 | 4.41 | 4.41 | Identical | On | On | Same | ¦12.030V 4.410A0¦ | ¦12.030V 4.410A ¦
805 | 3932.5843 | 15871.9101 | 15.84 | 15.84 | Identical | 3932.5843 | 4.41 | 4.42 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 4.410A0¦ | ¦12.030V 4.420A ¦
806 | 3937.4695 | 15891.6268 | 15.86 | 15.86 | Identical | 3937.4695 | 4.42 | 4.42 | Identical | On | On | Same | ¦12.030V 4.420A0¦ | ¦12.030V 4.420A ¦
807 | 3942.3547 | 15911.3434 | 15.88 | 15.88 | Identical | 3942.3547 | 4.43 | 4.43 | Identical | On | On | Same | ¦12.030V 4.430A0¦ | ¦12.030V 4.430A ¦
808 | 3947.2399 | 15931.0601 | 15.9 | 15.89 | Within 10 mV (calibration line) | 3947.2399 | 4.43 | 4.43 | Identical | On | On | Same | ¦12.030V 4.430A0¦ | ¦12.030V 4.430A ¦
809 | 3952.1251 | 15950.7767 | 15.92 | 15.91 | Within 10 mV (calibration line) | 3952.1251 | 4.44 | 4.44 | Identical | On | On | Same | ¦12.030V 4.440A0¦ | ¦12.030V 4.440A ¦
810 | 3957.0103 | 15970.4934 | 15.94 | 15.93 | Within 10 mV (calibration line) | 3957.0103 | 4.44 | 4.44 | Identical | On | On | Same | ¦12.030V 4.440A0¦ | ¦12.030V 4.440A ¦
811 | 3961.8955 | 15990.2101 | 15.96 | 15.95 | Within 10 mV (calibration line) | 3961.8955 | 4.45 | 4.45 | Identical | On | On | Same | ¦12.030V 4.450A0¦ | ¦12.030V 4.450A ¦
812 | 3966.7807 | 16009.9267 | 15.97 | 15.97 | Identical | 3966.7807 | 4.45 | 4.45 | Identical | On | On | Same | ¦12.030V 4.450A0¦ | ¦12.030V 4.450A ¦
813 | 3971.6659 | 16029.6434 | 15.99 | 15.99 | Identical | 3971.6659 | 4.46 | 4.46 | Identical | On | On | Same | ¦12.030V 4.460A0¦ | ¦12.030V 4.460A ¦
814 | 3976.5511 | 16049.36 | 16 | 16.01 | Within 10 mV (calibration line) | 3976.5511 | 4.46 | 4.46 | Identical | On | On | Same | ¦12.030V 4.460A0¦ | ¦12.030V 4.460A ¦
815 | 3981.4362 | 16069.0767 | 16.02 | 16.03 | Within 10 mV (calibration line) | 3981.4362 | 4.47 | 4.47 | Identical | On | On | Same | ¦12.030V 4.470A0¦ | ¦12.030V 4.470A ¦
816 | 3986.3214 | 16088.7934 | 16.04 | 16.05 | Within 10 mV (calibration line) | 3986.3214 | 4.47 | 4.47 | Identical | On | On | Same | ¦12.030V 4.470A0¦ | ¦12.030V 4.470A ¦
817 | 3991.2066 | 16108.51 | 16.06 | 16.07 | Within 10 mV (calibration line) | 3991.2066 | 4.48 | 4.48 | Identical | On | On | Same | ¦12.030V 4.480A0¦ | ¦12.030V 4.480A ¦
818 | 3996.0918 | 16128.2267 | 16.08 | 16.09 | Within 10 mV (calibration line) | 3996.0918 | 4.49 | 4.49 | Identical | On | On | Same | ¦12.030V 4.490A0¦ | ¦12.030V 4.490A ¦
819 | 4000.977 | 16147.9433 | 16.1 | 16.11 | Within 10 mV (calibration line) | 4000.977 | 4.49 | 4.49 | Identical | On | On | Same | ¦12.030V 4.490A0¦ | ¦12.030V 4.490A ¦
820 | 4005.8622 | 16167.66 | 16.12 | 16.13 | Within 10 mV (calibration line) | 4005.8622 | 4.5 | 4.5 | Identical | On | On | Same | ¦12.030V 4.500A0¦ | ¦12.030V 4.500A ¦
821 | 4010.7474 | 16187.3766 | 16.14 | 16.15 | Within 10 mV (calibration line) | 4010.7474 | 4.5 | 4.5 | Identical | On | On | Same | ¦12.030V 4.500A0¦ | ¦12.030V 4.500A ¦
822 | 4015.6326 | 16207.0933 | 16.16 | 16.17 | Within 10 mV (calibration line) | 4015.6326 | 4.51 | 4.51 | Identical | On | On | Same | ¦12.030V 4.510A0¦ | ¦12.030V 4.510A ¦
823 | 4020.5178 | 16226.81 | 16.18 | 16.19 | Within 10 mV (calibration line) | 4020.5178 | 4.51 | 4.51 | Identical | On | On | Same | ¦12.030V 4.510A0¦ | ¦12.030V 4.510A ¦
824 | 4025.403 | 16246.5266 | 16.2 | 16.21 | Within 10 mV (calibration line) | 4025.403 | 4.52 | 4.52 | Identical | On | On | Same | ¦12.030V 4.520A0¦ | ¦12.030V 4.520A ¦
825 | 4030.2882 | 16266.2433 | 16.22 | 16.23 | Within 10 mV (calibration line) | 4030.2882 | 4.52 | 4.52 | Identical | On | On | Same | ¦12.030V 4.520A0¦ | ¦12.030V 4.520A ¦
826 | 4035.1734 | 16285.9599 | 16.24 | 16.25 | Within 10 mV (calibration line) | 4035.1734 | 4.53 | 4.53 | Identical | On | On | Same | ¦12.030V 4.530A0¦ | ¦12.030V 4.530A ¦
827 | 4040.0586 | 16305.6766 | 16.26 | 16.27 | Within 10 mV (calibration line) | 4040.0586 | 4.54 | 4.54 | Identical | On | On | Same | ¦12.030V 4.540A0¦ | ¦12.030V 4.540A ¦
828 | 4044.9438 | 16325.3933 | 16.28 | 16.29 | Within 10 mV (calibration line) | 4044.9438 | 4.54 | 4.54 | Identical | On | On | Same | ¦12.030V 4.540A0¦ | ¦12.030V 4.540A ¦
829 | 4049.829 | 16345.1099 | 16.3 | 16.31 | Within 10 mV (calibration line) | 4049.829 | 4.55 | 4.55 | Identical | On | On | Same | ¦12.030V 4.550A0¦ | ¦12.030V 4.550A ¦
830 | 4054.7142 | 16364.8266 | 16.32 | 16.33 | Within 10 mV (calibration line) | 4054.7142 | 4.55 | 4.55 | Identical | On | On | Same | ¦12.030V 4.550A0¦ | ¦12.030V 4.550A ¦
831 | 4059.5994 | 16384.5432 | 16.34 | 16.35 | Within 10 mV (calibration line) | 4059.5994 | 4.56 | 4.56 | Identical | On | On | Same | ¦12.030V 4.560A0¦ | ¦12.030V 4.560A ¦
832 | 4064.4846 | 16404.2599 | 16.36 | 16.37 | Within 10 mV (calibration line) | 4064.4846 | 4.56 | 4.56 | Identical | On | On | Same | ¦12.030V 4.560A0¦ | ¦12.030V 4.560A ¦
833 | 4069.3698 | 16423.9766 | 16.38 | 16.39 | Within 10 mV (calibration line) | 4069.3698 | 4.57 | 4.57 | Identical | On | On | Same | ¦12.030V 4.570A0¦ | ¦12.030V 4.570A ¦
834 | 4074.255 | 16443.6932 | 16.4 | 16.41 | Within 10 mV (calibration line) | 4074.255 | 4.57 | 4.57 | Identical | On | On | Same | ¦12.030V 4.570A0¦ | ¦12.030V 4.570A ¦
835 | 4079.1402 | 16463.4099 | 16.42 | 16.42 | Identical | 4079.1402 | 4.58 | 4.58 | Identical | On | On | Same | ¦12.030V 4.580A0¦ | ¦12.030V 4.580A ¦
836 | 4084.0254 | 16483.1265 | 16.44 | 16.44 | Identical | 4084.0254 | 4.59 | 4.59 | Identical | On | On | Same | ¦12.030V 4.590A0¦ | ¦12.030V 4.590A ¦
837 | 4088.9106 | 16502.8432 | 16.46 | 16.46 | Identical | 4088.9106 | 4.59 | 4.59 | Identical | On | On | Same | ¦12.030V 4.590A0¦ | ¦12.030V 4.590A ¦
838 | 4093.7958 | 16522.5598 | 16.48 | 16.48 | Identical | 4093.7958 | 4.6 | 4.6 | Identical | On | On | Same | ¦12.030V 4.600A0¦ | ¦12.030V 4.600A ¦
839 | 4098.681 | 16542.2765 | 16.5 | 16.5 | Identical | 4098.681 | 4.6 | 4.6 | Identical | On | On | Same | ¦12.030V 4.600A0¦ | ¦12.030V 4.600A ¦
840 | 4103.5662 | 16561.9932 | 16.52 | 16.52 | Identical | 4103.5662 | 4.61 | 4.61 | Identical | On | On | Same | ¦12.030V 4.610A0¦ | ¦12.030V 4.610A ¦
841 | 4108.4514 | 16581.7098 | 16.54 | 16.54 | Identical | 4108.4514 | 4.61 | 4.61 | Identical | On | On | Same | ¦12.030V 4.610A0¦ | ¦12.030V 4.610A ¦
842 | 4113.3366 | 16601.4265 | 16.56 | 16.56 | Identical | 4113.3366 | 4.62 | 4.62 | Identical | On | On | Same | ¦12.030V 4.620A0¦ | ¦12.030V 4.620A ¦
843 | 4118.2218 | 16621.1431 | 16.58 | 16.58 | Identical | 4118.2218 | 4.62 | 4.62 | Identical | On | On | Same | ¦12.030V 4.620A0¦ | ¦12.030V 4.620A ¦
844 | 4123.107 | 16640.8598 | 16.6 | 16.6 | Identical | 4123.107 | 4.63 | 4.63 | Identical | On | On | Same | ¦12.030V 4.630A0¦ | ¦12.030V 4.630A ¦
845 | 4127.9922 | 16660.5765 | 16.62 | 16.62 | Identical | 4127.9922 | 4.63 | 4.64 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 4.630A0¦ | ¦12.030V 4.640A ¦
846 | 4132.8774 | 16680.2931 | 16.64 | 16.64 | Identical | 4132.8774 | 4.64 | 4.64 | Identical | On | On | Same | ¦12.030V 4.640A0¦ | ¦12.030V 4.640A ¦
847 | 4137.7626 | 16700.0098 | 16.66 | 16.66 | Identical | 4137.7626 | 4.65 | 4.65 | Identical | On | On | Same | ¦12.030V 4.650A0¦ | ¦12.030V 4.650A ¦
848 | 4142.6478 | 16719.7264 | 16.67 | 16.68 | Within 10 mV (calibration line) | 4142.6478 | 4.65 | 4.65 | Identical | On | On | Same | ¦12.030V 4.650A0¦ | ¦12.030V 4.650A ¦
849 | 4147.533 | 16739.4431 | 16.69 | 16.7 | Within 10 mV (calibration line) | 4147.533 | 4.66 | 4.66 | Identical | On | On | Same | ¦12.030V 4.660A0¦ | ¦12.030V 4.660A ¦
850 | 4152.4182 | 16759.1597 | 16.71 | 16.72 | Within 10 mV (calibration line) | 4152.4182 | 4.66 | 4.66 | Identical | On | On | Same | ¦12.030V 4.660A0¦ | ¦12.030V 4.660A ¦
851 | 4157.3034 | 16778.8764 | 16.73 | 16.74 | Within 10 mV (calibration line) | 4157.3034 | 4.67 | 4.67 | Identical | On | On | Same | ¦12.030V 4.670A0¦ | ¦12.030V 4.670A ¦
852 | 4162.1886 | 16798.5931 | 16.75 | 16.76 | Within 10 mV (calibration line) | 4162.1886 | 4.67 | 4.67 | Identical | On | On | Same | ¦12.030V 4.670A0¦ | ¦12.030V 4.670A ¦
853 | 4167.0738 | 16818.3097 | 16.77 | 16.78 | Within 10 mV (calibration line) | 4167.0738 | 4.68 | 4.68 | Identical | On | On | Same | ¦12.030V 4.680A0¦ | ¦12.030V 4.680A ¦
854 | 4171.959 | 16838.0264 | 16.79 | 16.8 | Within 10 mV (calibration line) | 4171.959 | 4.68 | 4.68 | Identical | On | On | Same | ¦12.030V 4.680A0¦ | ¦12.030V 4.680A ¦
855 | 4176.8442 | 16857.743 | 16.81 | 16.82 | Within 10 mV (calibration line) | 4176.8442 | 4.69 | 4.69 | Identical | On | On | Same | ¦12.030V 4.690A0¦ | ¦12.030V 4.690A ¦
856 | 4181.7294 | 16877.4597 | 16.83 | 16.84 | Within 10 mV (calibration line) | 4181.7294 | 4.69 | 4.7 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 4.690A0¦ | ¦12.030V 4.700A ¦
857 | 4186.6146 | 16897.1764 | 16.85 | 16.86 | Within 10 mV (calibration line) | 4186.6146 | 4.7 | 4.7 | Identical | On | On | Same | ¦12.030V 4.700A0¦ | ¦12.030V 4.700A ¦
858 | 4191.4998 | 16916.893 | 16.87 | 16.88 | Within 10 mV (calibration line) | 4191.4998 | 4.71 | 4.71 | Identical | On | On | Same | ¦12.030V 4.710A0¦ | ¦12.030V 4.710A ¦
859 | 4196.385 | 16936.6097 | 16.89 | 16.9 | Within 10 mV (calibration line) | 4196.385 | 4.71 | 4.71 | Identical | On | On | Same | ¦12.030V 4.710A0¦ | ¦12.030V 4.710A ¦
860 | 4201.2702 | 16956.3263 | 16.91 | 16.92 | Within 10 mV (calibration line) | 4201.2702 | 4.72 | 4.72 | Identical | On | On | Same | ¦12.030V 4.720A0¦ | ¦12.030V 4.720A ¦
861 | 4206.1553 | 16976.043 | 16.93 | 16.94 | Within 10 mV (calibration line) | 4206.1553 | 4.72 | 4.72 | Identical | On | On | Same | ¦12.030V 4.720A0¦ | ¦12.030V 4.720A ¦
862 | 4211.0405 | 16995.7596 | 16.95 | 16.96 | Within 10 mV (calibration line) | 4211.0405 | 4.73 | 4.73 | Identical | On | On | Same | ¦12.030V 4.730A0¦ | ¦12.030V 4.730A ¦
863 | 4215.9257 | 17015.4763 | 16.97 | 16.97 | Identical | 4215.9257 | 4.73 | 4.73 | Identical | On | On | Same | ¦12.030V 4.730A0¦ | ¦12.030V 4.730A ¦
864 | 4220.8109 | 17035.193 | 16.99 | 16.99 | Identical | 4220.8109 | 4.74 | 4.74 | Identical | On | On | Same | ¦12.030V 4.740A0¦ | ¦12.030V 4.740A ¦
865 | 4225.6961 | 17054.9096 | 17.01 | 17.01 | Identical | 4225.6961 | 4.74 | 4.74 | Identical | On | On | Same | ¦12.030V 4.740A0¦ | ¦12.030V 4.740A ¦
866 | 4230.5813 | 17074.6263 | 17.03 | 17.03 | Identical | 4230.5813 | 4.75 | 4.75 | Identical | On | On | Same | ¦12.030V 4.750A0¦ | ¦12.030V 4.750A ¦
867 | 4235.4665 | 17094.3429 | 17.05 | 17.05 | Identical | 4235.4665 | 4.75 | 4.75 | Identical | On | On | Same | ¦12.030V 4.750A0¦ | ¦12.030V 4.750A ¦
868 | 4240.3517 | 17114.0596 | 17.07 | 17.07 | Identical | 4240.3517 | 4.76 | 4.76 | Identical | On | On | Same | ¦12.030V 4.760A0¦ | ¦12.030V 4.760A ¦
869 | 4245.2369 | 17133.7763 | 17.09 | 17.09 | Identical | 4245.2369 | 4.77 | 4.77 | Identical | On | On | Same | ¦12.030V 4.770A0¦ | ¦12.030V 4.770A ¦
870 | 4250.1221 | 17153.4929 | 17.11 | 17.11 | Identical | 4250.1221 | 4.77 | 4.77 | Identical | On | On | Same | ¦12.030V 4.770A0¦ | ¦12.030V 4.770A ¦
871 | 4255.0073 | 17173.2096 | 17.13 | 17.13 | Identical | 4255.0073 | 4.78 | 4.78 | Identical | On | On | Same | ¦12.030V 4.780A0¦ | ¦12.030V 4.780A ¦
872 | 4259.8925 | 17192.9262 | 17.15 | 17.15 | Identical | 4259.8925 | 4.78 | 4.78 | Identical | On | On | Same | ¦12.030V 4.780A0¦ | ¦12.030V 4.780A ¦
873 | 4264.7777 | 17212.6429 | 17.17 | 17.17 | Identical | 4264.7777 | 4.79 | 4.79 | Identical | On | On | Same | ¦12.030V 4.790A0¦ | ¦12.030V 4.790A ¦
874 | 4269.6629 | 17232.3596 | 17.19 | 17.19 | Identical | 4269.6629 | 4.79 | 4.79 | Identical | On | On | Same | ¦12.030V 4.790A0¦ | ¦12.030V 4.790A ¦
875 | 4274.5481 | 17252.0762 | 17.21 | 17.21 | Identical | 4274.5481 | 4.8 | 4.8 | Identical | On | On | Same | ¦12.030V 4.800A0¦ | ¦12.030V 4.800A ¦
876 | 4279.4333 | 17271.7929 | 17.23 | 17.23 | Identical | 4279.4333 | 4.8 | 4.8 | Identical | On | On | Same | ¦12.030V 4.800A0¦ | ¦12.030V 4.800A ¦
877 | 4284.3185 | 17291.5095 | 17.25 | 17.25 | Identical | 4284.3185 | 4.81 | 4.81 | Identical | On | On | Same | ¦12.030V 4.810A0¦ | ¦12.030V 4.810A ¦
878 | 4289.2037 | 17311.2262 | 17.27 | 17.27 | Identical | 4289.2037 | 4.82 | 4.82 | Identical | On | On | Same | ¦12.030V 4.820A0¦ | ¦12.030V 4.820A ¦
879 | 4294.0889 | 17330.9428 | 17.29 | 17.29 | Identical | 4294.0889 | 4.82 | 4.82 | Identical | On | On | Same | ¦12.030V 4.820A0¦ | ¦12.030V 4.820A ¦
880 | 4298.9741 | 17350.6595 | 17.31 | 17.31 | Identical | 4298.9741 | 4.83 | 4.83 | Identical | On | On | Same | ¦12.030V 4.830A0¦ | ¦12.030V 4.830A ¦
881 | 4303.8593 | 17370.3762 | 17.33 | 17.33 | Identical | 4303.8593 | 4.83 | 4.83 | Identical | On | On | Same | ¦12.030V 4.830A0¦ | ¦12.030V 4.830A ¦
882 | 4308.7445 | 17390.0928 | 17.35 | 17.35 | Identical | 4308.7445 | 4.84 | 4.84 | Identical | On | On | Same | ¦12.030V 4.840A0¦ | ¦12.030V 4.840A ¦
883 | 4313.6297 | 17409.8095 | 17.36 | 17.37 | Within 10 mV (calibration line) | 4313.6297 | 4.84 | 4.84 | Identical | On | On | Same | ¦12.030V 4.840A0¦ | ¦12.030V 4.840A ¦
884 | 4318.5149 | 17429.5261 | 17.38 | 17.39 | Within 10 mV (calibration line) | 4318.5149 | 4.85 | 4.85 | Identical | On | On | Same | ¦12.030V 4.850A0¦ | ¦12.030V 4.850A ¦
885 | 4323.4001 | 17449.2428 | 17.4 | 17.41 | Within 10 mV (calibration line) | 4323.4001 | 4.85 | 4.85 | Identical | On | On | Same | ¦12.030V 4.850A0¦ | ¦12.030V 4.850A ¦
886 | 4328.2853 | 17468.9595 | 17.42 | 17.43 | Within 10 mV (calibration line) | 4328.2853 | 4.86 | 4.86 | Identical | On | On | Same | ¦12.030V 4.860A0¦ | ¦12.030V 4.860A ¦
887 | 4333.1705 | 17488.6761 | 17.44 | 17.45 | Within 10 mV (calibration line) | 4333.1705 | 4.86 | 4.86 | Identical | On | On | Same | ¦12.030V 4.860A0¦ | ¦12.030V 4.860A ¦
888 | 4338.0557 | 17508.3928 | 17.46 | 17.47 | Within 10 mV (calibration line) | 4338.0557 | 4.87 | 4.87 | Identical | On | On | Same | ¦12.030V 4.870A0¦ | ¦12.030V 4.870A ¦
889 | 4342.9409 | 17528.1094 | 17.48 | 17.49 | Within 10 mV (calibration line) | 4342.9409 | 4.87 | 4.88 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 4.870A0¦ | ¦12.030V 4.880A ¦
890 | 4347.8261 | 17547.8261 | 17.5 | 17.5 | Identical | 4347.8261 | 4.88 | 4.88 | Identical | On | On | Same | ¦12.030V 4.880A0¦ | ¦12.030V 4.880A ¦
891 | 4352.7113 | 17567.5427 | 17.52 | 17.52 | Identical | 4352.7113 | 4.89 | 4.89 | Identical | On | On | Same | ¦12.030V 4.890A0¦ | ¦12.030V 4.890A ¦
892 | 4357.5965 | 17587.2594 | 17.54 | 17.54 | Identical | 4357.5965 | 4.89 | 4.89 | Identical | On | On | Same | ¦12.030V 4.890A0¦ | ¦12.030V 4.890A ¦
893 | 4362.4817 | 17606.9761 | 17.56 | 17.56 | Identical | 4362.4817 | 4.9 | 4.9 | Identical | On | On | Same | ¦12.030V 4.900A0¦ | ¦12.030V 4.900A ¦
894 | 4367.3669 | 17626.6927 | 17.58 | 17.58 | Identical | 4367.3669 | 4.9 | 4.9 | Identical | On | On | Same | ¦12.030V 4.900A0¦ | ¦12.030V 4.900A ¦
895 | 4372.2521 | 17646.4094 | 17.6 | 17.6 | Identical | 4372.2521 | 4.91 | 4.91 | Identical | On | On | Same | ¦12.030V 4.910A0¦ | ¦12.030V 4.910A ¦
896 | 4377.1373 | 17666.126 | 17.62 | 17.62 | Identical | 4377.1373 | 4.91 | 4.91 | Identical | On | On | Same | ¦12.030V 4.910A0¦ | ¦12.030V 4.910A ¦
897 | 4382.0225 | 17685.8427 | 17.64 | 17.64 | Identical | 4382.0225 | 4.92 | 4.92 | Identical | On | On | Same | ¦12.030V 4.920A0¦ | ¦12.030V 4.920A ¦
898 | 4386.9077 | 17705.5594 | 17.66 | 17.66 | Identical | 4386.9077 | 4.92 | 4.92 | Identical | On | On | Same | ¦12.030V 4.920A0¦ | ¦12.030V 4.920A ¦
899 | 4391.7929 | 17725.276 | 17.68 | 17.68 | Identical | 4391.7929 | 4.93 | 4.93 | Identical | On | On | Same | ¦12.030V 4.930A0¦ | ¦12.030V 4.930A ¦
900 | 4396.6781 | 17744.9927 | 17.7 | 17.7 | Identical | 4396.6781 | 4.93 | 4.94 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 4.930A0¦ | ¦12.030V 4.940A ¦
901 | 4401.5633 | 17764.7093 | 17.72 | 17.72 | Identical | 4401.5633 | 4.94 | 4.94 | Identical | On | On | Same | ¦12.030V 4.940A0¦ | ¦12.030V 4.940A ¦
902 | 4406.4485 | 17784.426 | 17.74 | 17.74 | Identical | 4406.4485 | 4.95 | 4.95 | Identical | On | On | Same | ¦12.030V 4.950A0¦ | ¦12.030V 4.950A ¦
903 | 4411.3337 | 17804.1426 | 17.76 | 17.76 | Identical | 4411.3337 | 4.95 | 4.96 | +0.01 A (nearest step) | On | On | Same | ¦12.030V 4.950A0¦ | ¦12.030V 4.960A ¦
904 | 4416.2189 | 17823.8593 | 17.78 | 17.78 | Identical | 4416.2189 | 4.96 | 4.96 | Identical | On | On | Same | ¦12.030V 4.960A0¦ | ¦12.030V 4.960A ¦
905 | 4421.1041 | 17843.576 | 17.8 | 17.8 | Identical | 4421.1041 | 4.97 | 4.97 | Identical | On | On | Same | ¦12.030V 4.970A0¦ | ¦12.030V 4.970A ¦
906 | 4425.9893 | 17863.2926 | 17.82 | 17.82 | Identical | 4425.9893 | 4.97 | 4.97 | Identical | On | On | Same | ¦12.030V 4.970A0¦ | ¦12.030V 4.970A ¦
907 | 4430.8745 | 17883.0093 | 17.84 | 17.84 | Identical | 4430.8745 | 4.98 | 4.98 | Identical | On | On | Same | ¦12.030V 4.980A0¦ | ¦12.030V 4.980A ¦
908 | 4435.7596 | 17902.7259 | 17.86 | 17.86 | Identical | 4435.7596 | 4.98 | 4.98 | Identical | On | On | Same | ¦12.030V 4.980A0¦ | ¦12.030V 4.980A ¦
909 | 4440.6448 | 17922.4426 | 17.88 | 17.88 | Identical | 4440.6448 | 4.99 | 4.99 | Identical | On | On | Same | ¦12.030V 4.990A0¦ | ¦12.030V 4.990A ¦
910 | 4445.53 | 17942.1593 | 17.9 | 17.9 | Identical | 4445.53 | 4.99 | 4.99 | Identical | On | On | Same | ¦12.030V 4.990A0¦ | ¦12.030V 4.990A ¦
911 | 4450.4152 | 17961.8759 | 17.92 | 17.92 | Identical | 4450.4152 | 5 | 5 | Identical | On | On | Same | ¦12.030V 5.000A0¦ | ¦12.030V 5.000A ¦
912 | 4455.3004 | 17981.5926 | 17.94 | 17.94 | Identical | 4455.3004 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
913 | 4460.1856 | 18001.3092 | 17.96 | 17.96 | Identical | 4460.1856 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
914 | 4465.0708 | 18021.0259 | 17.98 | 17.98 | Identical | 4465.0708 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
915 | 4469.956 | 18040.7426 | 18 | 18 | Identical | 4469.956 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
916 | 4474.8412 | 18060.4592 | 18.02 | 18.02 | Identical | 4474.8412 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
917 | 4479.7264 | 18080.1759 | 18.04 | 18.03 | Within 10 mV (calibration line) | 4479.7264 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
918 | 4484.6116 | 18099.8925 | 18.05 | 18.05 | Identical | 4484.6116 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
919 | 4489.4968 | 18119.6092 | 18.07 | 18.07 | Identical | 4489.4968 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
920 | 4494.382 | 18139.3258 | 18.09 | 18.09 | Identical | 4494.382 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
921 | 4499.2672 | 18159.0425 | 18.11 | 18.11 | Identical | 4499.2672 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
922 | 4504.1524 | 18178.7592 | 18.13 | 18.13 | Identical | 4504.1524 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
923 | 4509.0376 | 18198.4758 | 18.15 | 18.15 | Identical | 4509.0376 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
924 | 4513.9228 | 18218.1925 | 18.17 | 18.17 | Identical | 4513.9228 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
925 | 4518.808 | 18237.9091 | 18.19 | 18.19 | Identical | 4518.808 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
926 | 4523.6932 | 18257.6258 | 18.21 | 18.21 | Identical | 4523.6932 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
927 | 4528.5784 | 18277.3425 | 18.23 | 18.23 | Identical | 4528.5784 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
928 | 4533.4636 | 18297.0591 | 18.25 | 18.25 | Identical | 4533.4636 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
929 | 4538.3488 | 18316.7758 | 18.27 | 18.27 | Identical | 4538.3488 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
930 | 4543.234 | 18336.4924 | 18.29 | 18.29 | Identical | 4543.234 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
931 | 4548.1192 | 18356.2091 | 18.31 | 18.31 | Identical | 4548.1192 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
932 | 4553.0044 | 18375.9257 | 18.33 | 18.33 | Identical | 4553.0044 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
933 | 4557.8896 | 18395.6424 | 18.35 | 18.35 | Identical | 4557.8896 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
934 | 4562.7748 | 18415.3591 | 18.37 | 18.37 | Identical | 4562.7748 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
935 | 4567.66 | 18435.0757 | 18.39 | 18.39 | Identical | 4567.66 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
936 | 4572.5452 | 18454.7924 | 18.41 | 18.41 | Identical | 4572.5452 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
937 | 4577.4304 | 18474.509 | 18.43 | 18.43 | Identical | 4577.4304 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
938 | 4582.3156 | 18494.2257 | 18.45 | 18.45 | Identical | 4582.3156 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
939 | 4587.2008 | 18513.9424 | 18.47 | 18.47 | Identical | 4587.2008 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
940 | 4592.086 | 18533.659 | 18.49 | 18.49 | Identical | 4592.086 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
941 | 4596.9712 | 18553.3757 | 18.51 | 18.51 | Identical | 4596.9712 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
942 | 4601.8564 | 18573.0923 | 18.53 | 18.53 | Identical | 4601.8564 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
943 | 4606.7416 | 18592.809 | 18.55 | 18.55 | Identical | 4606.7416 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
944 | 4611.6268 | 18612.5256 | 18.57 | 18.56 | Within 10 mV (calibration line) | 4611.6268 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
945 | 4616.512 | 18632.2423 | 18.59 | 18.58 | Within 10 mV (calibration line) | 4616.512 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
946 | 4621.3972 | 18651.959 | 18.61 | 18.6 | Within 10 mV (calibration line) | 4621.3972 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
947 | 4626.2824 | 18671.6756 | 18.63 | 18.62 | Within 10 mV (calibration line) | 4626.2824 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
948 | 4631.1676 | 18691.3923 | 18.65 | 18.64 | Within 10 mV (calibration line) | 4631.1676 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
949 | 4636.0528 | 18711.1089 | 18.67 | 18.66 | Within 10 mV (calibration line) | 4636.0528 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
950 | 4640.938 | 18730.8256 | 18.69 | 18.68 | Within 10 mV (calibration line) | 4640.938 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
951 | 4645.8232 | 18750.5423 | 18.71 | 18.7 | Within 10 mV (calibration line) | 4645.8232 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
952 | 4650.7084 | 18770.2589 | 18.73 | 18.72 | Within 10 mV (calibration line) | 4650.7084 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
953 | 4655.5936 | 18789.9756 | 18.74 | 18.74 | Identical | 4655.5936 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
954 | 4660.4787 | 18809.6922 | 18.76 | 18.76 | Identical | 4660.4787 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
955 | 4665.3639 | 18829.4089 | 18.78 | 18.78 | Identical | 4665.3639 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
956 | 4670.2491 | 18849.1255 | 18.8 | 18.8 | Identical | 4670.2491 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
957 | 4675.1343 | 18868.8422 | 18.82 | 18.82 | Identical | 4675.1343 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
958 | 4680.0195 | 18888.5589 | 18.84 | 18.84 | Identical | 4680.0195 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
959 | 4684.9047 | 18908.2755 | 18.86 | 18.86 | Identical | 4684.9047 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
960 | 4689.7899 | 18927.9922 | 18.88 | 18.88 | Identical | 4689.7899 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
961 | 4694.6751 | 18947.7088 | 18.9 | 18.9 | Identical | 4694.6751 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
962 | 4699.5603 | 18967.4255 | 18.92 | 18.92 | Identical | 4699.5603 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
963 | 4704.4455 | 18987.1422 | 18.94 | 18.94 | Identical | 4704.4455 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
964 | 4709.3307 | 19006.8588 | 18.96 | 18.96 | Identical | 4709.3307 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
965 | 4714.2159 | 19026.5755 | 18.98 | 18.98 | Identical | 4714.2159 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
966 | 4719.1011 | 19046.2921 | 19 | 19 | Identical | 4719.1011 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
967 | 4723.9863 | 19066.0088 | 19.02 | 19.02 | Identical | 4723.9863 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
968 | 4728.8715 | 19085.7255 | 19.04 | 19.04 | Identical | 4728.8715 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
969 | 4733.7567 | 19105.4421 | 19.06 | 19.06 | Identical | 4733.7567 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
970 | 4738.6419 | 19125.1588 | 19.08 | 19.08 | Identical | 4738.6419 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
971 | 4743.5271 | 19144.8754 | 19.1 | 19.1 | Identical | 4743.5271 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
972 | 4748.4123 | 19164.5921 | 19.12 | 19.11 | Within 10 mV (calibration line) | 4748.4123 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
973 | 4753.2975 | 19184.3087 | 19.14 | 19.13 | Within 10 mV (calibration line) | 4753.2975 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
974 | 4758.1827 | 19204.0254 | 19.16 | 19.15 | Within 10 mV (calibration line) | 4758.1827 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
975 | 4763.0679 | 19223.7421 | 19.18 | 19.17 | Within 10 mV (calibration line) | 4763.0679 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
976 | 4767.9531 | 19243.4587 | 19.2 | 19.19 | Within 10 mV (calibration line) | 4767.9531 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
977 | 4772.8383 | 19263.1754 | 19.22 | 19.21 | Within 10 mV (calibration line) | 4772.8383 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
978 | 4777.7235 | 19282.892 | 19.24 | 19.23 | Within 10 mV (calibration line) | 4777.7235 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
979 | 4782.6087 | 19302.6087 | 19.26 | 19.25 | Within 10 mV (calibration line) | 4782.6087 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
980 | 4787.4939 | 19322.3254 | 19.28 | 19.27 | Within 10 mV (calibration line) | 4787.4939 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
981 | 4792.3791 | 19342.042 | 19.3 | 19.29 | Within 10 mV (calibration line) | 4792.3791 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
982 | 4797.2643 | 19361.7587 | 19.32 | 19.31 | Within 10 mV (calibration line) | 4797.2643 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
983 | 4802.1495 | 19381.4753 | 19.34 | 19.33 | Within 10 mV (calibration line) | 4802.1495 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
984 | 4807.0347 | 19401.192 | 19.36 | 19.35 | Within 10 mV (calibration line) | 4807.0347 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
985 | 4811.9199 | 19420.9086 | 19.38 | 19.37 | Within 10 mV (calibration line) | 4811.9199 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
986 | 4816.8051 | 19440.6253 | 19.4 | 19.39 | Within 10 mV (calibration line) | 4816.8051 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
987 | 4821.6903 | 19460.342 | 19.42 | 19.41 | Within 10 mV (calibration line) | 4821.6903 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
988 | 4826.5755 | 19480.0586 | 19.44 | 19.43 | Within 10 mV (calibration line) | 4826.5755 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
989 | 4831.4607 | 19499.7753 | 19.45 | 19.45 | Identical | 4831.4607 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
990 | 4836.3459 | 19519.4919 | 19.47 | 19.47 | Identical | 4836.3459 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
991 | 4841.2311 | 19539.2086 | 19.49 | 19.49 | Identical | 4841.2311 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
992 | 4846.1163 | 19558.9253 | 19.51 | 19.51 | Identical | 4846.1163 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
993 | 4851.0015 | 19578.6419 | 19.53 | 19.53 | Identical | 4851.0015 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
994 | 4855.8867 | 19598.3586 | 19.55 | 19.55 | Identical | 4855.8867 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
995 | 4860.7719 | 19618.0752 | 19.57 | 19.57 | Identical | 4860.7719 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
996 | 4865.6571 | 19637.7919 | 19.59 | 19.59 | Identical | 4865.6571 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
997 | 4870.5423 | 19657.5085 | 19.61 | 19.61 | Identical | 4870.5423 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
998 | 4875.4275 | 19677.2252 | 19.63 | 19.63 | Identical | 4875.4275 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
999 | 4880.3127 | 19696.9419 | 19.65 | 19.64 | Within 10 mV (calibration line) | 4880.3127 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
1000 | 4885.1979 | 19716.6585 | 19.67 | 19.66 | Within 10 mV (calibration line) | 4885.1979 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
1001 | 4890.083 | 19736.3752 | 19.69 | 19.68 | Within 10 mV (calibration line) | 4890.083 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
1002 | 4894.9682 | 19756.0918 | 19.71 | 19.7 | Within 10 mV (calibration line) | 4894.9682 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
1003 | 4899.8534 | 19775.8085 | 19.73 | 19.72 | Within 10 mV (calibration line) | 4899.8534 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
1004 | 4904.7386 | 19795.5252 | 19.75 | 19.74 | Within 10 mV (calibration line) | 4904.7386 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
1005 | 4909.6238 | 19815.2418 | 19.77 | 19.76 | Within 10 mV (calibration line) | 4909.6238 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
1006 | 4914.509 | 19834.9585 | 19.79 | 19.78 | Within 10 mV (calibration line) | 4914.509 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
1007 | 4919.3942 | 19854.6751 | 19.81 | 19.8 | Within 10 mV (calibration line) | 4919.3942 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
1008 | 4924.2794 | 19874.3918 | 19.83 | 19.82 | Within 10 mV (calibration line) | 4924.2794 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
1009 | 4929.1646 | 19894.1085 | 19.85 | 19.84 | Within 10 mV (calibration line) | 4929.1646 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
1010 | 4934.0498 | 19913.8251 | 19.87 | 19.86 | Within 10 mV (calibration line) | 4934.0498 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
1011 | 4938.935 | 19933.5418 | 19.89 | 19.88 | Within 10 mV (calibration line) | 4938.935 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
1012 | 4943.8202 | 19953.2584 | 19.91 | 19.9 | Within 10 mV (calibration line) | 4943.8202 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
1013 | 4948.7054 | 19972.9751 | 19.93 | 19.92 | Within 10 mV (calibration line) | 4948.7054 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
1014 | 4953.5906 | 19992.6917 | 19.95 | 19.94 | Within 10 mV (calibration line) | 4953.5906 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
1015 | 4958.4758 | 20012.4084 | 19.97 | 19.96 | Within 10 mV (calibration line) | 4958.4758 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
1016 | 4963.361 | 20032.1251 | 19.99 | 19.98 | Within 10 mV (calibration line) | 4963.361 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
1017 | 4968.2462 | 20051.8417 | 20.01 | 20 | Within 10 mV (calibration line) | 4968.2462 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
1018 | 4973.1314 | 20071.5584 | 20.03 | 20.02 | Within 10 mV (calibration line) | 4973.1314 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
1019 | 4978.0166 | 20091.275 | 20.05 | 20.04 | Within 10 mV (calibration line) | 4978.0166 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
1020 | 4982.9018 | 20110.9917 | 20.07 | 20.06 | Within 10 mV (calibration line) | 4982.9018 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
1021 | 4987.787 | 20130.7084 | 20.09 | 20.08 | Within 10 mV (calibration line) | 4987.787 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
1022 | 4992.6722 | 20150.425 | 20.11 | 20.1 | Within 10 mV (calibration line) | 4992.6722 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦
1023 | 4997.5574 | 20170.1417 | 20.13 | OL | ADC full scale: now OL | 4997.5574 | ERR (999.23) | OL | Above 5.00 A: now OL | Off | On | Different | ¦12.030V 999.2A0¦ | ¦12.030V OL ¦

## Stability Tests
Display stability and resolution tests
Each noise scenario is 300 measurement cycles (about 2.8 minutes of operation). Changes = number of times the displayed value changed.
Scenario | Stimulus | Cycles | v0.3.0 changes | v1.0.0 changes, outlier rejection off (limit 1023) | v1.0.0 changes | v0.3.0 values shown (count) | v1.0.0 values shown, outlier rejection off | v1.0.0 values shown
Voltage on a 10 mV boundary | ADC 611.40 (12.035 V after calibration, a v1.0.0 step boundary), uniform ±1 count noise | 300 | 158 | 3 | 3 | 12.01: 18, 12.03: 160, 12.05: 122 | 12.03: 132, 12.04: 168 | 12.03: 132, 12.04: 168
Voltage inside a 10 mV step | ADC 611.15 (12.030 V), uniform ±1 count noise | 300 | 188 | 0 | 0 | 12.01: 49, 12.03: 151, 12.05: 100 | 12.03: 300 | 12.03: 300
Current on the 0.71/0.72 A boundary | ADC 130.70 (638.5 mV), uniform ±1 count noise | 300 | 142 | 1 | 1 | 0.71: 121, 0.72: 179 | 0.71: 89, 0.72: 211 | 0.71: 89, 0.72: 211
Current 0.72 A inside a step | ADC 131.62 (643 mV), uniform ±1 count noise | 300 | 30 | 0 | 0 | 0.72: 285, 0.73: 15 | 0.72: 300 | 0.72: 300
Current 0.72 A with sample dropouts (original symptom) | ADC 131.62, ±0.5 count noise, 2% of samples replaced by ADC 4 (0.02 A) | 300 | 10 | 215 | 0 | 0.02: 5, 0.72: 295 | 0.65: 1, 0.66: 1, 0.67: 14, 0.68: 15, 0.69: 10, 0.7: 77, 0.71: 111, 0.72: 71 | 0.72: 300
Totals |  |  | 528 | 219 | 4 | Display changes avoided by v1.0.0 vs v0.3.0: | 99.24%
Slow ramps (input changes by 0.1 ADC count per cycle, then back down)
Ramp | Direction | Cycles | Ideal changes | v0.3.0 changes | v1.0.0 changes | v0.3.0 max error vs ideal | v1.0.0 max error vs ideal | Comment
I 0.49-0.78 A (codes 100-160) | up | 601 | 34 | 33 | 33 | 0.01 | 0.01 | Both follow the input; v1.0.0 holds a step only inside the 1.5 mV hysteresis band.
I 0.49-0.78 A (codes 100-160) | down | 601 | 34 | 33 | 33 | 0.01 | 0.01
V 3.9-5.1 V (codes 200-260) | up | 601 | 118 | 59 | 117 | 0.01 | 0.01 | v0.3.0 skips every other 10 mV value (1 ADC count = 19.7 mV). v1.0.0 shows every value (3.960-5.130 V after calibration), never against the input: the 4.150 V band-edge reversal is gone.
V 3.9-5.1 V (codes 200-260) | down | 601 | 118 | 59 | 117 | 0.01 | 0.01
V 0.71-0.95 V low range (codes 36-48) | up | 241 | 19 | 9 | 19 | 0.79 | 0.01 | v0.3.0 can never show 0.820 or 0.830 and runs backwards at 0.800 V. v1.0.0 shows 0.000 up to 0.790 V after calibration, then every 10 mV value in order. Ideal changes now count the calibrated values.
V 0.71-0.95 V low range (codes 36-48) | down | 241 | 19 | 9 | 19 | 0.79 | 0.01
Outlier rejection: behavior outside the tested firmware scenarios (Python model, 4,000 trials each, same algorithm)
Signal | Description |  | Plain average bias (counts) | Plain average spread (counts) | Outlier rejection bias (counts) | Outlier rejection spread (counts) | Better choice
Normal noise ±2 counts | Typical for this board |  | 0.002 | 0.145 | 0.002 | 0.145 | Equal
Noise ±8 counts | At the rejection limit |  | 0.023 | 0.57 | 0.069 | 0.837 | Plain average (slightly)
Noise ±20 counts, symmetric | Much noisier than expected |  | -0.022 | 1.445 | 0.302 | 3.094 | Plain average
Symmetric ripple ±15 counts | Large real ripple at the ADC |  | 0.004 | 1.083 | 0.234 | 2.438 | Plain average
Asymmetric ripple: 10% of the time 30 counts low | Real dips that are part of the true average |  | 0.007 | 1.135 | 3 | 0.077 | Plain average
2% dropouts to 0 | Isolated bad samples (original symptom) |  | -10.083 | 8.905 | 0.002 | 0.076 | Outlier rejection
1 count = 4.9 mV at the ADC pin = 5.4 mA of current = 19.7 mV of channel voltage. Decide with bench test BT-43.

## LCD Scenarios
LCD, LED and screen-clear behavior in event scenarios
Rows are exactly what the 16x2 LCD shows (top row = CH2, bottom row = CH1). LEDs = D3 (CH1) then D5 (CH2), 1 = on. Clears = lcd.clear() calls in the cycle.
Scenario | Step | Stimulus | v0.3.0 top row | v0.3.0 bottom row | v0.3.0 LEDs | v0.3.0 clears | v1.0.0 top row | v1.0.0 bottom row | v1.0.0 LEDs | v1.0.0 clears | Comment
Over range | 0 | CH1 12.03 V / 0.72 A, CH2 12.03 V / 0.72 A | ¦12.030V 0.720A0¦ | ¦12.030V 0.720A0¦ | 11 | 0 | ¦12.030V 0.720A ¦ | ¦12.030V 0.720A ¦ | 11 | 0
Over range | 1 | CH1 current ADC 1000 (above 5.00 A) | ¦12.030V 0.720A0¦ | ¦12.030V 999.2A0¦ | 01 | 0 | ¦12.030V 0.720A ¦ | ¦12.030V OL ¦ | 11 | 0 | v0.3.0: garbage "999.2A0", LED off. v1.0.0: OL, LED on
Over range | 2 | CH1 current back to 0.72 A | ¦12.030V 0.720A0¦ | ¦12.030V 0.720A0¦ | 11 | 0 | ¦12.030V 0.720A ¦ | ¦12.030V 0.720A ¦ | 11 | 0 | v0.3.0 keeps a stale "0" in column 15
Over range | 3 | CH1 voltage ADC 1023 (full scale) | ¦12.030V 0.720A0¦ | ¦20.130V 0.720A0¦ | 11 | 0 | ¦12.030V 0.720A ¦ | ¦ OL 0.720A ¦ | 11 | 0 | v0.3.0 shows 20.130V for an unknown voltage
Over range | 4 | CH1 voltage back to 12.03 V | ¦12.030V 0.720A0¦ | ¦12.030V 0.720A0¦ | 11 | 0 | ¦12.030V 0.720A ¦ | ¦12.030V 0.720A ¦ | 11 | 0
Over range | 5 | CH1 voltage and current both over range | ¦12.030V 0.720A0¦ | ¦20.130V 999.2A0¦ | 01 | 0 | ¦12.030V 0.720A ¦ | ¦ OL OL ¦ | 11 | 0
Over range | 6 | Both back to normal | ¦12.030V 0.720A0¦ | ¦12.030V 0.720A0¦ | 11 | 0 | ¦12.030V 0.720A ¦ | ¦12.030V 0.720A ¦ | 11 | 0
Over range | 7 | CH1 current ADC 911 (5.00 A, last table value) | ¦12.030V 0.720A0¦ | ¦12.030V 5.000A0¦ | 11 | 0 | ¦12.030V 0.720A ¦ | ¦12.030V 5.000A ¦ | 11 | 0
Over range | 8 | CH1 current ADC 912 (first code above 5.00 A) | ¦12.030V 0.720A0¦ | ¦12.030V 999.2A0¦ | 01 | 0 | ¦12.030V 0.720A ¦ | ¦12.030V OL ¦ | 11 | 0
LED | 0 | CH1 current ADC 0 (0.00 A); CH2 ADC 254 / 0 A | ¦5.0100V 0.000A0¦ | ¦12.030V 0.000A0¦ | 00 | 0 | ¦ 5.020V 0.000A ¦ | ¦12.030V 0.000A ¦ | 00 | 0
LED | 1 | CH1 current ADC 1 (0.00 A (0.005 A, rounds to 0)); CH2 ADC 254 / 0 A | ¦5.0100V 0.000A0¦ | ¦12.030V 0.000A0¦ | 00 | 0 | ¦ 5.020V 0.000A ¦ | ¦12.030V 0.000A ¦ | 00 | 0
LED | 2 | CH1 current ADC 2 (0.01 A); CH2 ADC 254 / 0 A | ¦5.0100V 0.000A0¦ | ¦12.030V 0.010A0¦ | 10 | 0 | ¦ 5.020V 0.000A ¦ | ¦12.030V 0.010A ¦ | 10 | 0
LED | 3 | CH1 current ADC 3 (0.01 A); CH2 ADC 254 / 0 A | ¦5.0100V 0.000A0¦ | ¦12.030V 0.010A0¦ | 10 | 0 | ¦ 5.020V 0.000A ¦ | ¦12.030V 0.010A ¦ | 10 | 0
LED | 4 | CH1 current ADC 911 (5.00 A); CH2 ADC 254 / 0 A | ¦5.0100V 0.000A0¦ | ¦12.030V 5.000A0¦ | 10 | 0 | ¦ 5.020V 0.000A ¦ | ¦12.030V 5.000A ¦ | 10 | 0
LED | 5 | CH1 current ADC 912 (above 5.00 A); CH2 ADC 254 / 0 A | ¦5.0100V 0.000A0¦ | ¦12.030V 999.2A0¦ | 00 | 0 | ¦ 5.020V 0.000A ¦ | ¦12.030V OL ¦ | 10 | 0 | v0.3.0 LED off above 5.00 A; v1.0.0 LED on
LED | 6 | CH1 current ADC 1023 (ADC full scale); CH2 ADC 254 / 0 A | ¦5.0100V 0.000A0¦ | ¦12.030V 999.2A0¦ | 00 | 0 | ¦ 5.020V 0.000A ¦ | ¦12.030V OL ¦ | 10 | 0 | v0.3.0 LED off above 5.00 A; v1.0.0 LED on
LED | 7 | CH1 current ADC 0 (0.00 A); CH2 ADC 254 / 0 A | ¦5.0100V 0.000A0¦ | ¦12.030V 0.000A0¦ | 00 | 0 | ¦ 5.020V 0.000A ¦ | ¦12.030V 0.000A ¦ | 00 | 0
Temperature | 0 | Temperature pin 0.645 V (ADC 132) | ¦5.0100V 0.000A0¦ | ¦12.030V 0.720A0¦ | 10 | 0 | ¦ 5.020V 0.000A ¦ | ¦12.030V 0.720A ¦ | 10 | 0
Temperature | 1 | Temperature pin 0.645 V (ADC 132) | ¦5.0100V 0.000A0¦ | ¦12.030V 0.720A0¦ | 10 | 0 | ¦ 5.020V 0.000A ¦ | ¦12.030V 0.720A ¦ | 10 | 0
Temperature | 2 | Temperature pin 0.684 V (ADC 140) | ¦ IC6-TEMP ¦ | ¦ ALERT ¦ | 10 | 1 | ¦ VREF-TEMP ¦ | ¦ ALERT ¦ | 10 | 1 | Alert on in both
Temperature | 3 | Temperature pin 0.684 V (ADC 140) | ¦ IC6-TEMP ¦ | ¦ ALERT ¦ | 10 | 1 | ¦ VREF-TEMP ¦ | ¦ ALERT ¦ | 10 | 0
Temperature | 4 | Temperature pin 0.679 V (ADC 139) | ¦5.0100V-T0.000A ¦ | ¦12.030VLE0.720A ¦ | 10 | 0 | ¦ VREF-TEMP ¦ | ¦ ALERT ¦ | 10 | 0 | v0.3.0 alert off at 0.679 V and leftover "-T" / "LE"; v1.0.0 stays in alert (hysteresis)
Temperature | 5 | Temperature pin 0.674 V (ADC 138) | ¦5.0100V-T0.000A ¦ | ¦12.030VLE0.720A ¦ | 10 | 0 | ¦ VREF-TEMP ¦ | ¦ ALERT ¦ | 10 | 0
Temperature | 6 | Temperature pin 0.679 V (ADC 139) | ¦5.0100V-T0.000A ¦ | ¦12.030VLE0.720A ¦ | 10 | 0 | ¦ VREF-TEMP ¦ | ¦ ALERT ¦ | 10 | 0
Temperature | 7 | Temperature pin 0.684 V (ADC 140) | ¦ IC6-TEMP ¦ | ¦ ALERT ¦ | 10 | 1 | ¦ VREF-TEMP ¦ | ¦ ALERT ¦ | 10 | 0
Temperature | 8 | Temperature pin 0.674 V (ADC 138) | ¦5.0100V-T0.000A ¦ | ¦12.030VLE0.720A ¦ | 10 | 0 | ¦ VREF-TEMP ¦ | ¦ ALERT ¦ | 10 | 0
Temperature | 9 | Temperature pin 0.669 V (ADC 137) | ¦5.0100V-T0.000A ¦ | ¦12.030VLE0.720A ¦ | 10 | 0 | ¦ 5.020V 0.000A ¦ | ¦12.030V 0.720A ¦ | 10 | 1 | v1.0.0 alert off below 0.670 V; screen cleared once, clean rows
Temperature | 10 | Temperature pin 0.664 V (ADC 136) | ¦5.0100V-T0.000A ¦ | ¦12.030VLE0.720A ¦ | 10 | 0 | ¦ 5.020V 0.000A ¦ | ¦12.030V 0.720A ¦ | 10 | 0
Temperature | 11 | Temperature pin 0.645 V (ADC 132) | ¦5.0100V-T0.000A ¦ | ¦12.030VLE0.720A ¦ | 10 | 0 | ¦ 5.020V 0.000A ¦ | ¦12.030V 0.720A ¦ | 10 | 0
Temperature | 12 | Temperature pin 0.645 V (ADC 132) | ¦5.0100V-T0.000A ¦ | ¦12.030VLE0.720A ¦ | 10 | 0 | ¦ 5.020V 0.000A ¦ | ¦12.030V 0.720A ¦ | 10 | 0
Temperature | 13 | Temperature pin 0.684 V (ADC 140) | ¦ IC6-TEMP ¦ | ¦ ALERT ¦ | 10 | 1 | ¦ VREF-TEMP ¦ | ¦ ALERT ¦ | 10 | 1 | Hovering 140/139: v0.3.0 toggles every cycle; v1.0.0 stays in alert
Temperature | 14 | Temperature pin 0.679 V (ADC 139) | ¦5.0100V-T0.000A ¦ | ¦12.030VLE0.720A ¦ | 10 | 0 | ¦ VREF-TEMP ¦ | ¦ ALERT ¦ | 10 | 0 | Hovering 140/139: v0.3.0 toggles every cycle; v1.0.0 stays in alert
Temperature | 15 | Temperature pin 0.684 V (ADC 140) | ¦ IC6-TEMP ¦ | ¦ ALERT ¦ | 10 | 1 | ¦ VREF-TEMP ¦ | ¦ ALERT ¦ | 10 | 0 | Hovering 140/139: v0.3.0 toggles every cycle; v1.0.0 stays in alert
Temperature | 16 | Temperature pin 0.679 V (ADC 139) | ¦5.0100V-T0.000A ¦ | ¦12.030VLE0.720A ¦ | 10 | 0 | ¦ VREF-TEMP ¦ | ¦ ALERT ¦ | 10 | 0 | Hovering 140/139: v0.3.0 toggles every cycle; v1.0.0 stays in alert
Temperature | 17 | Temperature pin 0.645 V (ADC 132) | ¦5.0100V-T0.000A ¦ | ¦12.030VLE0.720A ¦ | 10 | 0 | ¦ 5.020V 0.000A ¦ | ¦12.030V 0.720A ¦ | 10 | 1
Crosstalk | 0 | Model: first 0 read(s) after a channel switch return the previous channel value; CH1 17.70 V / 0.72 A | ¦5.0100V-T0.000A ¦ | ¦17.700VLE0.720A ¦ | 10 | 0 | ¦ 5.020V 0.000A ¦ | ¦17.700V 0.720A ¦ | 10 | 0
Crosstalk | 1 | Model: first 1 read(s) after a channel switch return the previous channel value; CH1 17.70 V / 0.72 A | ¦5.0100V-T0.000A ¦ | ¦17.700VLE0.720A ¦ | 10 | 0 | ¦ 5.020V 0.000A ¦ | ¦17.700V 0.720A ¦ | 10 | 0
Crosstalk | 2 | Model: first 2 read(s) after a channel switch return the previous channel value; CH1 17.70 V / 0.72 A | ¦0.0000V-T0.000A ¦ | ¦0.0000VLE0.000A ¦ | 00 | 0 | ¦ 5.020V 0.000A ¦ | ¦17.700V 0.720A ¦ | 10 | 0 | v0.3.0 wrong: it discards only 1 read
Note: the v0.3.0 rows in the Crosstalk block still carry the "-T" / "LE" leftovers from the temperature test before it; v0.3.0 never clears them.

## Findings
Findings and observations for v1.0.0 (open and closed)
Severity: Observation / Risk / Info / Cosmetic. None blocks the release; bench checks are in the Bench Test Plan.
ID | Severity | Status | Finding | Details | Evidence | Action
F-01 | Observation | Open - check on bench | Outlier rejection trade-off | v1.0.0 ignores samples more than 8 counts from the median. With 2% single-sample dropouts it shows 0.720 A in all 300 cycles; with rejection off (OUTLIER_LIMIT_COUNTS = 1023, a plain average) the display changed 215 times (0.65-0.72 A). v0.3.0 showed 0.02 A in 5 cycles. If the real signal has large asymmetric ripple (more than ±8 counts), the plain average is more accurate. | Tests A-07 to A-09; Stability Tests sheet | Keep 8 unless bench test BT-43 shows the plain average (1023) reads closer to the reference on the load that caused the oscillation.
F-02 | Observation | Fixed in final file | Display ran backwards at 0.800 V (closed) | v0.3.0 and the first v1.0.0 build: step 0 (0.800 V) got no +0.030 V offset, so a rising input showed 0.790, 0.810, 0.800, 0.840. The final file replaces the offset bands with one calibration line, so the display always rises: 0.800, 0.810, 0.820 ... | V-04, V-09, U-10; ramp RL; fine sweep | None. BT-03 / BT-13 still check accuracy at 0.80 V.
F-03 | Observation | Fixed in final file | Calibration band edges (closed) | The offset bands changed by 10-20 mV at 4.15, 5.98, 8.761, 13.01, 15.02 and 16.03 V, so the display could step back (4.170, 4.180, 4.170). The calibration line has no edges: 0 reversals in 40,922 ramp steps. | V-05, V-09; ramp RV; fine sweep | None. BT-04 / BT-14 still check accuracy at 4.15 V.
F-04 | Observation | Open - recommendation | No hysteresis on the OL limits | With ±1 count of noise exactly at full scale, the display alternates between OL and 20.100/20.110V (voltage) or 5.000A and OL (current), about every other cycle. | Tests E-03, E-04 | Only at 20.1 V or 5.00 A, above the supply rating (0.8-20 V, 4 A), so only in a fault. Optional: add a small hysteresis to the OL decision.
F-05 | Observation | Open - recommendation | LEDs keep their last state during a temperature alert | Measurements pause, so a LED may stay on (or off) until the alert ends. Same as v0.3.0. | Test L-04 | Optional: turn both LEDs off when the alert starts.
F-06 | Info | Open - recommendation | "WATCHDOG RESET" message may not appear on boards with Optiboot | Optiboot clears MCUSR before the firmware starts, so the reset cause reads 0. The watchdog itself still works (verified). | Test W-01 (no bootloader in simulation) | Optional: also read the reset cause Optiboot leaves in register r2. Not implemented because it cannot be tested here.
F-07 | Risk | Open - check on bench | Old Nano bootloader and the watchdog | Boards that only accept "ATmega328P (Old Bootloader)" do not disable the watchdog after a watchdog reset and can reset endlessly. | Known ATmegaBOOT behavior | Before flashing, check the board uploads with Processor "ATmega328P" (BT-01). Burn Optiboot otherwise.
F-08 | Info | Open - check on bench | Pull-ups on A4/A5 and D2/D4/D6 | The firmware enables pull-ups on pins it does not use. | Pin setup | Confirm on the PCB that none of these pins is connected (BT-02).
F-09 | Info | Open - recommendation | fabs() promotes float to double | Reported only with -Wdouble-promotion (lines 492, 533). On the ATmega328P double is the same 32-bit type as float, so results and speed are unchanged. | Test B-06 | No change needed; the Arduino IDE never enables this warning. fabsf() does not help: avr-libc defines it as fabs(). __builtin_fabsf() removes the warning with an identical binary (tested), but is less readable.
F-10 | Cosmetic | Open - recommendation | Two names for the current-sense voltage | The MAX9919 output in mV is called currentSenseMv in computeCurrent() but voltageSenseMv in findNearestCurrentStep(), which receives it. "voltageSense" can also be confused with the voltage channel. | Code review (lines 223, 501-515) | Optional: rename the parameter of findNearestCurrentStep() to currentSenseMv (6 places). A name-only change: the compiled program does not change.
F-11 | Cosmetic | Open - recommendation | Comment style | Line 457: "//return true or false" has no space after //, unlike every other comment. Line 561: example "5.030", but dtostrf(…, 6, 3) gives " 5.030". Line 584: debug example "vAdc=254.00 … V=5.030"; with the calibration line ADC 254 prints V=5.020. | Code review | Optional: change line 457 to "// Returns true or false" (or remove it), write " 5.030" on line 561, and V=5.020 on line 584. Comments only.
F-12 | Info | Open - recommendation | No version history in the source header | The file header does not list the firmware versions. | Code review | Optional: add a short history (v0.3.0, v1.0.0) to the header, or keep this workbook with the release.
F-13 | Info | Open - check on bench | Voltage accuracy above 4 V not yet proven | The calibration line is a fit of the v0.3.0 offset table. The Fluke 15B+ check agrees at every point, but above 4 V the Fluke tolerance (±0.5 % + 3 digits, 10 mV resolution: ±55 mV at 5 V, ±130 mV at 20 V) is larger than the meter's estimated ±(0.1 % + 2 digits). Below 4 V the meter is within -7 to -2 mV. | V-12; Voltage Accuracy sheet | Check 3-4 points between 5 and 20 V against a reference of 0.02 % or better (BT-49). If errors exceed the spec, recalculate VOLTAGE_CAL_GAIN / VOLTAGE_CAL_OFFSET_MV from those points.
F-14 | Info | Open - recommendation | Highest displayed voltage is 20.110 V | The ADC saturates at 20.170 V (measured). OL starts at ADC 1022.5 (20.160 V measured), which the calibration shows as 20.110 V. OL above 20.40 V would need a larger input divider (ratio >= 4.082): a hardware change. | V-02, V-11 | None for a 0.8-20 V rating. State the range as 0.80-20.00 V in the spec.

## Memory & Timing
Memory and timing, measured
Flash and static RAM from avr-size; stack and cycle time measured on the real firmware in simulation.
Nano usable flash (bytes, 2 KB bootloader reserved) |  |  |  | 30720
ATmega328P RAM (bytes) |  |  |  | 2048
Watchdog timeout (ms) |  |  |  | 2000
Metric | v0.3.0 | v1.0.0 (production) | v1.0.0 (debug build) | Note
Flash (bytes) | 8052 | 9358 | 10358 | Final file: 410 B less than the 7-band build (9,768 B)
Flash used | 26.21% | 30.46% | 33.72% | Formula
Static RAM (bytes) | 1319 | 273 | 287 | v0.3.0 keeps the 1 KB lookup table in RAM
Peak stack (bytes) | 133 | 210 | 226 | v1.0.0 includes a 128 B sample buffer for outlier rejection
Free RAM at peak (bytes) | 596 | 1565 | 1535 | Formula
Measurement work per cycle (ms) | 13 | 42.7 | 47.2 | v0.3.0: 573 ms loop minus 560 ms of delays
Measurement period (ms) | 573 | 560 | 560 | v1.0.0 uses a millis() scheduler
Watchdog margin (timeout / work) | n/a | 46.8384 | 42.3729 | Formula; >1 means safe
ADC conversions per cycle | 10 | 282 | 282
Extra sort time, worst case, per 64-sample channel (ms) | n/a | 3.476 | 3.476 | Measured: reverse-ordered samples (4.668 ms) minus sorted (1.192 ms)
Worst-case measurement work per cycle (ms) | 13 | 56.604 | 61.104 | Formula: measured work + 4 channels x worst-case sort
Worst-case watchdog margin (timeout / work) | n/a | 35.3332 | 32.7311 | Formula

## Bench Test Plan
Bench test plan: fill in on the real hardware
Simulation cannot check the analog hardware. Use a calibrated reference meter.
How to use: type only in the yellow cells. Error and Result calculate by themselves. Row 9 is an example.
Current tolerance (A) |  |  |  | 0.01 | Your spec: ±0.01 A
Voltage tolerance (V) |  |  |  | 0.02 | Fixed part of the voltage spec (2 digits = 20 mV)
Voltage tolerance, % of reading |  |  |  | 0.1% | Estimated spec ±(0.1 % + 2 digits); see Voltage Accuracy
ID | Area | Channel | Setup / what to check | Target | Reference reading | Meter display | Error | Tolerance | Result | Notes
EX | Example | CH1 | Example row: 12 V setpoint, fixed resistor load | 12 | 12.012 | 12.03 | 0.018 | 0.032 | PASS | Example only
BT-01 | Bootloader | Board | Upload with Tools > Processor "ATmega328P" (not Old Bootloader). Type PASS if it uploads. | [6 empty] | F-07
BT-02 | PCB | Board | Continuity: D2, D4, D6, A4, A5 not connected to anything. Type PASS or FAIL. | [6 empty] | F-08
BT-03 | Voltage | CH1 | Output 0.80 V, light load | 0.8 |  |  |  | 0.021 |  | Former F-02 (closed): 0.800 V now in order
BT-04 | Voltage | CH1 | Output 4.15 V, light load | 4.15 |  |  |  | 0.024 |  | Former F-03 (closed): no band edge now
BT-05 | Voltage | CH1 | Output 1.00 V, light load | 1 |  |  |  | 0.021
BT-06 | Voltage | CH1 | Output 3.30 V, light load | 3.3 |  |  |  | 0.023
BT-07 | Voltage | CH1 | Output 5.00 V, light load | 5 |  |  |  | 0.025
BT-08 | Voltage | CH1 | Output 9.00 V, light load | 9 |  |  |  | 0.029
BT-09 | Voltage | CH1 | Output 12.00 V, light load | 12 |  |  |  | 0.032
BT-10 | Voltage | CH1 | Output 15.00 V, light load | 15 |  |  |  | 0.035
BT-11 | Voltage | CH1 | Output 18.00 V, light load | 18 |  |  |  | 0.038
BT-12 | Voltage | CH1 | Output 20.00 V, light load | 20 |  |  |  | 0.04
BT-13 | Voltage | CH2 | Output 0.80 V, light load | 0.8 |  |  |  | 0.021 |  | Former F-02 (closed): 0.800 V now in order
BT-14 | Voltage | CH2 | Output 4.15 V, light load | 4.15 |  |  |  | 0.024 |  | Former F-03 (closed): no band edge now
BT-15 | Voltage | CH2 | Output 1.00 V, light load | 1 |  |  |  | 0.021
BT-16 | Voltage | CH2 | Output 3.30 V, light load | 3.3 |  |  |  | 0.023
BT-17 | Voltage | CH2 | Output 5.00 V, light load | 5 |  |  |  | 0.025
BT-18 | Voltage | CH2 | Output 9.00 V, light load | 9 |  |  |  | 0.029
BT-19 | Voltage | CH2 | Output 12.00 V, light load | 12 |  |  |  | 0.032
BT-20 | Voltage | CH2 | Output 15.00 V, light load | 15 |  |  |  | 0.035
BT-21 | Voltage | CH2 | Output 18.00 V, light load | 18 |  |  |  | 0.038
BT-22 | Voltage | CH2 | Output 20.00 V, light load | 20 |  |  |  | 0.04
BT-23 | Current | CH1 | Load current 0.05 A | 0.05 |  |  |  | 0.01
BT-24 | Current | CH1 | Load current 0.10 A | 0.1 |  |  |  | 0.01
BT-25 | Current | CH1 | Load current 0.50 A | 0.5 |  |  |  | 0.01
BT-26 | Current | CH1 | Load current 0.72 A | 0.72 |  |  |  | 0.01
BT-27 | Current | CH1 | Load current 1.00 A | 1 |  |  |  | 0.01
BT-28 | Current | CH1 | Load current 2.49 A | 2.49 |  |  |  | 0.01 |  | Old gap zone
BT-29 | Current | CH1 | Load current 2.50 A | 2.5 |  |  |  | 0.01 |  | Old gap zone
BT-30 | Current | CH1 | Load current 3.00 A | 3 |  |  |  | 0.01
BT-31 | Current | CH1 | Load current 4.00 A | 4 |  |  |  | 0.01
BT-32 | Current | CH1 | Load current 4.90 A | 4.9 |  |  |  | 0.01
BT-33 | Current | CH2 | Load current 0.05 A | 0.05 |  |  |  | 0.01
BT-34 | Current | CH2 | Load current 0.10 A | 0.1 |  |  |  | 0.01
BT-35 | Current | CH2 | Load current 0.50 A | 0.5 |  |  |  | 0.01
BT-36 | Current | CH2 | Load current 0.72 A | 0.72 |  |  |  | 0.01
BT-37 | Current | CH2 | Load current 1.00 A | 1 |  |  |  | 0.01
BT-38 | Current | CH2 | Load current 2.49 A | 2.49 |  |  |  | 0.01 |  | Old gap zone
BT-39 | Current | CH2 | Load current 2.50 A | 2.5 |  |  |  | 0.01 |  | Old gap zone
BT-40 | Current | CH2 | Load current 3.00 A | 3 |  |  |  | 0.01
BT-41 | Current | CH2 | Load current 4.00 A | 4 |  |  |  | 0.01
BT-42 | Current | CH2 | Load current 4.90 A | 4.9 |  |  |  | 0.01
BT-43 | Stability | Both | Load that caused the oscillation, 60 s. Compare display with reference, then repeat with OUTLIER_LIMIT_COUNTS = 1023. Keep the setting closer to the reference. Type PASS or FAIL. | [6 empty] | F-01
BT-44 | Stability | Both | Fixed resistor load, 60 s: last digit should not flicker. Type PASS or FAIL.
BT-45 | LED | Both | LED off at 0.00 A, on with any load. Type PASS or FAIL.
BT-46 | Display | Both | Rows show "12.030V 0.720A" style, no stray characters after load changes. Type PASS or FAIL.
BT-47 | Temperature | Board | Optional: warm IC6 gently; alert appears, clears after cooling, rows clean. Type PASS or FAIL. | [6 empty] | F-05
BT-48 | Soak | Both | 24 h run with load: splash "WATCHDOG RESET" never appears, readings stable. Type PASS or FAIL.
BT-49 | Voltage | Both | With a reference of 0.02 % or better (bench DMM or voltage reference): 5, 10, 15 and 20 V. Every error within ±(0.1 % + 20 mV)? Type PASS or FAIL. | [6 empty] | F-13
Bench results
|  |  | Tests passed (excluding example) | 0
|  |  | Tests failed | 0
|  |  | Tests not done yet | 49
Note on E5: Assumed value, not from your spec. Replace with your voltage accuracy requirement.

## Voltage Accuracy
Voltage accuracy: bench check, error budget and comparison
Final firmware: displayed = measured x VOLTAGE_CAL_GAIN (0.99578) + VOLTAGE_CAL_OFFSET_MV (31.0), rounded to 10 mV. Yellow cells are inputs; everything else is calculated.
Estimated meter spec, % of reading | 0.1% | Estimate from the error budget below: ±(0.1 % + 2 digits)
Estimated meter spec, fixed part (V) | 0.02 | 2 digits of the 10 mV display
Fluke 15B+ DC accuracy, % of reading | 0.5% | Fluke 15B+ technical data: ±(0.5 % + 3 digits) on the 4 V and 40 V ranges
1. Bench check against a Fluke 15B+ (27 Sep 2026, final firmware; readings from the user)
Point | Supply display (V) | Fluke 15B+ (V) | Difference (mV) | Fluke resolution (V) | Fluke tolerance ±(mV) | Inside Fluke tolerance | Meter spec ±(mV) | Inside meter spec | Can the Fluke decide?
P1 | 0.8 | 0.807 | -7 | 0.001 | 7 | Yes | 20.8 | Yes | Yes: Fluke tighter than the spec
P2 | 1.01 | 1.016 | -6 | 0.001 | 8.1 | Yes | 21 | Yes | Yes: Fluke tighter than the spec
P3 | 2.01 | 2.012 | -2 | 0.001 | 13.1 | Yes | 22 | Yes | Yes: Fluke tighter than the spec
P4 | 5 | 4.99 | 10 | 0.01 | 55 | Yes | 25 | Yes | No: Fluke tolerance too wide
P5 | 8 | 8 | 0 | 0.01 | 70 | Yes | 28 | Yes | No: Fluke tolerance too wide
P6 | 10 | 9.98 | 20 | 0.01 | 79.9 | Yes | 30 | Yes | No: Fluke tolerance too wide
P7 | 12.01 | 11.98 | 30 | 0.01 | 89.9 | Yes | 32 | Yes | No: Fluke tolerance too wide
P8 | 15.01 | 14.96 | 50 | 0.01 | 104.8 | Yes | 35 | No | No: Fluke tolerance too wide
P9 | 18 | 17.96 | 40 | 0.01 | 119.8 | Yes | 38 | No | No: Fluke tolerance too wide
P10 | 20 | 19.95 | 50 | 0.01 | 129.8 | Yes | 40 | No | No: Fluke tolerance too wide
Points inside the Fluke tolerance | 10 of 10
Points where the Fluke can decide | 3
Largest ¦difference¦ where the Fluke can decide (mV) | #VALUE! | Below 4 V the Fluke uses 1 mV resolution: the meter is within a few mV there.
Largest ¦difference¦ overall (mV) | #VALUE! | Above 4 V every difference is inside the Fluke's own tolerance, so the Fluke cannot tell which instrument is right (F-13).
2. Error budget (estimated, at the display)
Source | Size ±(mV) | Fixable in firmware? | Note
Display rounding to 10 mV | 5 | No | Half of one 10 mV step
Hysteresis lag near a step | 3 | No (it stops flicker) | VOLTAGE_HYSTERESIS_MV, only right at a step boundary
Calibration fit | 10 | Only by recalibrating | Old table had 10 mV steps (±5 mV) + line residual (up to 5 mV)
ADC non-linearity (INL) | 10 | No | ATmega328P typical 0.5 count; 1 count = 19.7 mV at the input
Noise and drift | 5 | Already reduced | 64-sample average; MAX6250 reference
Typical total (root sum of squares, mV) | 16.1 | Sources rarely peak together: about ±15 mV typical.
Worst case (all errors add up, mV) | 33 | Covered by the ±(0.1 % + 2 digits) spec over most of the range.
3. Comparison with commercial bench supplies (voltage readback specification)
Supply | Class | % of reading | Fixed part (mV) | ±mV at 1 V | ±mV at 5 V | ±mV at 12 V | ±mV at 20 V | Ratio at 12 V vs this meter | Source
Siglent SPD3303X-E | Entry level | 0.5% | 20 | 25 | 45 | 80 | 120 | 2.5 | Siglent SPD3000X data sheet: ±(0.5 % + 2 digits), 10 mV resolution
This meter (estimate) | DIY, 10-bit ADC | 0.1% | 20 | 21 | 25 | 32 | 40 | 1 | Error budget above
Rigol DP832 | Mid-range lab | 0.05% | 10 | 10.5 | 12.5 | 16 | 20 | 0.5 | Rigol DP800 data sheet: ±(0.05 % + 10 mV), 25 °C ± 5 °C, 1 year
Siglent SPD3303X | Mid-range lab | 0.03% | 10 | 10.3 | 11.5 | 13.6 | 16 | 0.43 | Siglent SPD3000X data sheet: ±(0.03 % + 10 mV)
Keysight E36312A (25 V outputs) | Professional | 0.04% | 3 | 3.4 | 5 | 7.8 | 11 | 0.24 | Keysight E36300 data sheet: ±(0.04 % + 3 mV), 23 °C ± 5 °C, 1 year
Ratio below 1 = more accurate than this meter. Commercial figures are guaranteed specs for every unit; this meter's figure is an estimate for one unit, confirmed only below 4 V.
Sources
Fluke 15B+ technical data: https://mohanmeasurements.com/wp-content/uploads/2024/10/fluke-15b-plus-digital-multimeter.pdf
Rigol DP800 series data sheet: https://www.rigol.com/dam/global/downloads/brochures/en/data-sheet/dc-powers/DP800_DataSheet_EN.pdf
Siglent SPD3303X / X-E data sheet: https://www.batronix.com/pdf/Siglent/SPD3303X/SPD3303X-Datasheet.pdf
Keysight E36312A data sheet: https://docs.alltest.net/manual/Alltest-Agilent-Keysight-E36312A-Datasheet-29381-.pdf
ATmega328P data sheet, ADC characteristics (INL typical 0.5 LSB)

## Method
How the firmware was tested
Firmware under test | v0.3.0: 2_channel_dual_voltage_current_meter_fw_v0.3.0.ino. v1.0.0: 2_channel_dual_voltage_current_meter_fw_v1.0.0.ino (release file; compiles to exactly the binary every suite ran on).
Compiler | avr-gcc 7.3.0, Arduino AVR core 1.8.6, LiquidCrystal 1.0.7, target ATmega328P at 16 MHz (Arduino Nano, eightanaloginputs variant). Strict C++ without Arduino IDE auto-prototypes, -Wall -Wextra.
Simulator | simavr: cycle-accurate ATmega328P. Float math, timers, watchdog and UART behave as on the chip.
Functional test bench | Both sketches compiled unmodified. analogRead, analogReference, pinMode, digitalWrite, millis, delay and delayMicroseconds replaced by instrumented versions. The LCD is a model of the HD44780 memory built on the real Arduino Print class, so every print() formats exactly as on the real display.
Input models | Exact ADC codes (sweeps); dithered fractional codes (ramps); uniform random noise; random single-sample dropouts; channel crosstalk (first 1 or 2 reads after a switch return the previous channel value).
Timing and stack | Real firmware with the real ADC, real LCD library and real serial port; only main() replaced to time each cycle with micros(). Free RAM measured by filling unused RAM with a pattern before startup.
Watchdog test | Infinite loop injected inside analogRead(A2); the simulator performs a real watchdog reset and reboot.
Independent models | Python re-implementations (float32) of the v0.3.0 voltage algorithm, the v1.0.0 calibration-line voltage algorithm and a nearest-step current model, used to cross-check every ADC code, every unit case and a 51,151-point fine sweep.
Limits | The simulation proves the firmware logic, formatting and timing. It cannot check analog accuracy, real noise on your PCB, or the bootloader on your board: see Bench Test Plan.
Unit tests | The firmware functions are called directly on the simulated chip with scripted inputs (every step boundary, table midpoint, band edge and display value) and compared with independent float32 models.
Edge scenarios | Channel independence, large steps, noise at the OL limits, power-on with alert, alert during OL, watchdog register state, irregular loop() calls, and a 3,000-cycle random soak checked row by row.
Static checks | Extra compiler warnings, cppcheck static analysis, and negative tests that break each static_assert on purpose.
Totals | 2 x 1,024-code sweeps (voltage and current, both channels), 2,886 ramp cycles, 1,500 noise cycles, plus event scenarios: 6,473 result rows per version. Unit level: 234,322 function calls. Bench: 10-point voltage check against a Fluke 15B+.
Bench check | Real supply with the final firmware, 10 points 0.8-20 V, read with a Fluke 15B+ (±(0.5 % + 3 digits), Fluke 15B+ technical data). Evaluated on the Voltage Accuracy sheet.